## Main
| PMT# | Location | Warm Up Time | LED Setting | 700 V | 750 V | 775 V | 800 V | 825 V | 850 V | Drift Test @ 800 V | Fit Voltage | Calibrated | Calibrated Resolution | Adjusted Voltage | Calibrated 2 | Calibrated 2 Resolution | Fit Accuracy (%) | Gain Accuracy (%) | Drift (%) | PMT Resolution (%) | ABS Gain |  |  | Notes | [5 empty] | Old Data
S1T | LV2143 | S1T5 |  | 1 | 1327 | 1987 | 2365 | 2791 | 3261 | 3800 | 2739 | 758.1566 | 2057 | 283 |  |  |  | -2.09 | -2.09 | -1.9 | 13.758 | 2.09 |  | * | the reason fits usually underestimate the voltage needed to gain match and why the fit doesnot converge is likely due to the inclusion of 825 and 850 V measurement, where the relationship between voltage and gain is not the same exponential. The curve starts to level out.
|  |  |  | 2 | 9375 | 13167 | 15445 | 18005 | 20865 | 24082 | 17846 | 759.626 | 13883 | 319 | 760.8 | 13982 | 323 | -0.84 | -0.13 | -0.89 | 2.298 | 0.13
|  |  |  | 3 | 27229 | 37560 | 43775 | 50738 | 57759 | 63571 | 50816 | 778.4382 | 44719 | 396 | 779.6 | 45059 | 381 | -0.63 | 0.13 | 0.15 | 0.886 | 0.13
| LV2177 | S1T4 |  | 1 | 1929 | 2784 | 3300 | 3885 | 4533 | 5241 |  | 711.0938 | 2073 | 251 | 712 | 2087 | 249 | -1.3 | -0.62 |  | 12.108 |  |  |  | NEED TO INCREASE GAIN OF PMT |  |  |  | OLD DATA BEFORE RESISITOR CHANGE
|  |  |  | 2 | 13226 | 17951 | 20644 | 23539 | 26550 | 29654 |  | 708.5569 | 14028 | 454 |  |  |  | 0.2 | 0.2 |  | 3.236 |  |  |  | 3410 | 4545 | 5202 | 5918 | 6684 | 7509
|  |  |  | 3 | 33658 | 44009 | 49828 | 55932 | 61256 |  |  | 754.238 | 44904 | 363 |  |  |  | -0.21 | -0.21 |  | 0.808 |  |  |  | 10665 | 13825 | 15605 | 17511 | 19525 | 21671
| LV1762 | S1T3 |  | 1 | 1299 | 1944 | 2313 | 2737 | 3213 | 3749 | 2734 | 761.1007 | 2097 | 286 |  |  |  | -0.14 | -0.14 | -0.11 | 13.639 | 0.14
|  |  |  | 2 | 9769 | 13642 | 15939 | 18537 | 21416 | 24641 | 18509 | 753.361 | 13923 | 328 |  |  |  | -0.55 | -0.55 | -0.15 | 2.356 | 0.55
|  |  | 3:37PM->3:46PM | 3 | 29142 | 39313 | 45195 | 51658 | 57743 | 62903 | 51660 | 773.0777 | 44696 | 380 | 774.3 | 44986 | 384 | -0.68 | -0.03 | 0 | 0.85 | 0.03
| LV2085 | S1T2 |  | 1 | 1058 | 1666 | 2007 | 2391 | 2818 | 3289 | 2407 | 781.184 | 2117 | 221 |  |  |  | 0.8 | 0.8 | 0.66 | 10.439 | 0.8 | [8 empty] | 1205 | 1845 | 2225 | 2651 | 3122 | 3644
|  |  |  | 2 | 8238 | 11726 | 13857 | 16276 | 19032 | 22131 | 16408 | 776.5196 | 14108 | 281 |  |  |  | 0.77 | 0.77 | 0.8 | 1.992 | 0.77 | [8 empty] | 9020 | 12870 | 15207 | 17852 | 20862 | 24300
|  |  | 4:07PM->4:21PM | 3 | 25983 | 36884 | 43674 | 51504 | 59218 | Saturates | 51765 | 779.7565 | 45348 | 393 | 778.2 | 44881 | 394 | 0.77 | -0.27 | 0.5 | 0.867 | 0.27 | [8 empty] | 28554 | 40421 | 47996 | 56627 | 64887 | saturates
| LV2138 | S1T1 |  | 1 | 1553 | 2298 | 2745 | 3269 | 3841 | 4491 | 3284 | 737.4251 | 2094 | 231 |  |  |  | -0.29 | -0.29 | 0.46 | 11.032 | 0.29
|  |  |  | 2 | 10904 | 15497 | 18265 | 21306 | 24811 | 28823 | 21439 | 735.1453 | 14083 | 287 |  |  |  | 0.59 | 0.59 | 0.62 | 2.038 | 0.59
|  |  | 4:38PM->4:52PM | 3 | 32263 | 45005 | 52762 | 60851 | Saturates | Saturates | 61105 | 749.3545 | 45101 | 368 |  |  |  | 0.22 | 0.22 | 0.42 | 0.816 | 0.22
S1B | LV2125 | S1B5 | 7:39AM->8:52AM | 1 | 1300 | 1963 | 2375 | 2834 | 3348 | 3903 | 2924 | 758.1589 | 2088 | 312 |  |  |  | -0.57 | -0.57 | 3.08 | 14.943 | 0.57
|  |  |  | 2 | 9309 | 13262 | 15590 | 18223 | 21282 | 24590 | 18409 | 758.3424 | 13962 | 335 |  |  |  | -0.27 | -0.27 | 1.01 | 2.399 | 0.27
|  |  |  | 3 | 25883 | 35391 | 41429 | 48330 | 55919 | 62685 | 48519 | 788.0184 | 45013 | 415 |  |  |  | 0.03 | 0.03 | 0.39 | 0.922 | 0.03
| LV2148 | S1B4 | 7:39AM->9:20AM | 1 | 993 | 1422 | 1762 | 2120 | 2517 | 2962 | 2076 | 799.7763 | 2090 | 279 |  |  |  | -0.48 | -0.48 | -2.12 | 13.349 | 0.48
|  |  |  | 2 | 6938 | 9933 | 11703 | 13712 | 15962 | 18512 | 13722 | 803.5296 | 14022 | 316 |  |  |  | 0.16 | 0.16 | 0.07 | 2.254 | 0.16
|  |  |  | 3 | 19486 | 27302 | 31975 | 37137 | 42662 | 49234 | 37279 | 834.1423 | 44862 | 399 |  |  |  | -0.31 | -0.31 | 0.38 | 0.889 | 0.31
| LV2060 | S1B3 | 7:39AM->9:42AM | 1 | 1148 | 1763 | 2102 | 2486 | 2914 | 3408 | 2465 | 774.7804 | 2097 | 269 |  |  |  | -0.14 | -0.14 | -0.85 | 12.828 | 0.14
|  |  |  | 2 | 8217 | 11749 | 13860 | 16272 | 18988 | 22029 | 16147 | 776.3276 | 13933 | 306 |  |  |  | -0.48 | -0.48 | -0.77 | 2.196 | 0.48
|  |  |  | 3 | 23808 | 33559 | 39210 | 45424 | 52513 | 58678 | 45420 | 798.4455 | 44807 | 393 |  |  |  | -0.43 | -0.43 | -0.01 | 0.877 | 0.43
| LV2029 | S1B2 | 7:39AM->10:17AM | 1 | 1844 | 2721 | 3233 | 3850 | 4528 | 5240 | 3882 | 715.9915 | 2084 | 304 |  |  |  | -0.77 | -0.77 | 0.82 | 14.587 | 0.77
|  |  |  | 2 | 12849 | 17879 | 20827 | 24079 | 27610 | 31321 | 24004 | 712.3241 | 13888 | 333 |  |  |  | -0.81 | -0.81 | -0.31 | 2.398 | 0.81
|  |  |  | 3 | 33569 | 45086 | 51560 | 57521 | 62620 | saturates | 57642 | 751.5919 | 45439 | 372 | 750.5 | 45234 | 369 | 0.97 | 0.52 | 0.21 | 0.819 | 0.52
| LV2050 | S1B1 | 7:39AM->10:46AM | 1 | 1231 | 1858 | 2213 | 2614 | 3155 | 3662 | 2611 | 767.4795 | 2161 | 249 | 762.5 | 2094 | 245 | 2.82 | -0.29 | -0.11 | 11.522 | 0.29
|  |  |  | 2 | 8680 | 12427 | 14652 | 17305 | 20259 | 23697 | 17272 | 767.8394 | 13928 | 329 |  |  |  | -0.52 | -0.52 | -0.19 | 2.362 | 0.52
|  |  |  | 3 | 25413 | 35937 | 42498 | 50163 | 58065 | 64555 | 50127 | 783.5601 | 44726 | 410 |  |  |  | -0.61 | -0.61 | -0.07 | 0.917 | 0.61
S2T | LV2119 | S2T5 |  | 1 | 2234 | 3179 | 3754 | 4391 | 5098 | 5868 |  | 692.5197 | 2060 | 296 |  |  |  | -1.94 | -1.94 | #DIV/0! | 14.369 | 1.94
|  |  |  | 2 | 15053 | 20211 | 23099 | 26152 | 29627 | 33384 |  | 688.2432 | 13747 | 350 |  |  |  | -1.84 | -1.84 | #DIV/0! | 2.546 | 1.84
|  |  | 11:51AM->12:01PM | 3 | 41252 | 53964 | 60521 | Saturates | Saturates | Saturates |  | 714.9528 | 44395 | 398 |  |  |  | -1.36 | -1.36 | #VALUE! | 0.896 | 1.36 |  | * | the gain of LV2119 is too high. It was saturating and overflowing at 800 V with LED3
| LV2086 | S2T4 |  | 1 | 1029 | 1598 | 1923 | 2283 | 2693 | 3146 |  | 787.2003 | 2094 | 310 |  |  |  | -0.29 | -0.29 | #DIV/0! | 14.804 | 0.29
|  |  |  | 2 | 7965 | 11328 | 13365 | 15664 | 18245 | 21193 |  | 780.9475 | 13715 | 356 |  |  |  | -2.08 | -2.08 | #DIV/0! | 2.596 | 2.08
|  |  | 12:16PM-> | 3 | 25215 | 35583 | 41802 | 49079 | 56919 | 63442 |  | 785.3268 | 44384 | 439 |  |  |  | -1.39 | -1.39 | #DIV/0! | 0.989 | 1.39
| LV2193 | S2T3 | 2024-03-04 | 1 | out of range | 1112 | 1358 | 1614 | 1758 | 2132 | 1623 | 849 | 2137 | 295 |  |  |  | 1.73 | 1.73 | 0.55 | 13.804 | 1.73
|  |  |  | 2 | 6748 | 9231 | 10675 | 12279 | 14045 | 15990 | 12355 | 823 | 13961 | 354 |  |  |  | -0.28 | -0.28 | 0.62 | 2.536 | 0.28
|  |  | 7:37AM ->~ 8:00 | 3 | 21332 | 28345 | 32295 | 36526 | 41101 | 45967 | 36593 | 844.1 | 44880 | 403 |  |  |  | -0.27 | -0.27 | 0.18 | 0.898 | 0.27
| LV2182 | S2T2 |  | 1 | 1179 | 1785 | 2126 | 2517 | 2946 | 3429 | 2450 | 772.893 | 2034 | 313 |  |  |  | -3.24 | -3.24 | -2.73 | 15.388 | 3.24
|  |  |  | 2 | 8666 | 12193 | 14332 | 16723 | 19381 | 22383 | 16675 | 770.6308 | 13886 | 352 |  |  |  | -0.82 | -0.82 | -0.29 | 2.535 | 0.82
|  |  | 8:50 ->9:11AM | 3 | 26554 | 36418 | 42551 | 49506 | 56788 | 62925 | 49551 | 783.102 | 44762 | 422 | 784.5 | 45150 | 418 | -0.53 | 0.33 | 0.09 | 0.943 | 0.33
| LV2134 | S2T1 |  | 1 | 1396 | 1894 | 2172 | 2474 | 2801 | 3160 | 2388 | 768.7726 | 2023 | 275 | 770.5 | 2040 | 272 | -3.81 | -2.94 | -3.6 | 13.594 | 2.94
|  |  |  | 2 | 10810 | 14338 | 16332 | 18455 | 20744 | 23181 | 18169 | 745.3385 | 13751 | 325 | 746.5 | 13850 | 326 | -1.81 | -1.08 | -1.57 | 2.363 | 1.08
|  |  | 9:41AM->9:53AM | 3 | 33526 | 42849 | 47872 | 53195 | 58532 | 62886 | 52583 | 760.1468 | 44575 | 386 | 762.2 | 44973 | 394 | -0.95 | -0.06 | -1.16 | 0.866 | 0.06
S2B | LV2201 | S2B5 | 12:56PM->1:05PM | 1 | 1082 | 1711 | 2065 | 2476 | 2941 | 3455 | 2499 | 776.7 | 2114 | 243 |  |  |  | 0.66 | 0.66 | 0.92 | 11.495 | 0.66
|  |  |  | 2 | 8409 | 11963 | 14067 | 16463 | 19129 | 22089 | 16526 | 774 | 14036 | 298 |  |  |  | 0.26 | 0.26 | 0.38 | 2.123 | 0.26
|  |  |  | 3 | 25276 | 35007 | 40994 | 47754 | 55126 | 61556 | 47911 | 789.9 | 45094 | 392 |  |  |  | 0.21 | 0.21 | 0.33 | 0.869 | 0.21
| LV2032 | S2B4 | 1:20PM->1:30PM | 1 | 1238 | 1819 | 2123 | 2465 | 2831 | 3240 | 2472 | 772.7 | 2097 | 265 |  |  |  | -0.14 | -0.14 | 0.28 | 12.637 | 0.14
|  |  |  | 2 | 9414 | 12714 | 14585 | 16626 | 18869 | 21224 | 16671 | 766.9 | 13995 | 309 |  |  |  | -0.04 | -0.04 | 0.27 | 2.208 | 0.04
|  |  |  | 3 | 27836 | 36546 | 41440 | 46695 | 52252 | 57992 | 46781 | 791.8 | 45059 | 382 |  |  |  | 0.13 | 0.13 | 0.18 | 0.848 | 0.13
| LV2070 | S2B3 | 1:49PM->1:59PM | 1 | Too low | 1335 | 1615 | 1884 | 2181 | 2515 | 1889 | 817.5265 | 2097 | 233 |  |  |  | -0.14 | -0.14 | 0.26 | 11.111 | 0.14 |  | * | LV2070 was corrected on D10 instead of C10 because QDC2.11 does not work.
|  |  |  | 2 | 7428 | 10168 | 11769 | 13511 | 15416 | 17468 | 13489 | 806.8264 | 13975 | 287 |  |  |  | -0.18 | -0.18 | -0.16 | 2.054 | 0.18
|  |  |  | 3 | 22700 | 30183 | 34392 | 38877 | 43627 | 48722 | 38899 | 831.3762 | 44913 | 359 |  |  |  | -0.19 | -0.19 | 0.06 | 0.799 | 0.19
| LV1788 | S2B2 | 2:19PM->2:29PM | 1 | 1031 | 1546 | 1846 | 2177 | 2555 | 2970 | 2137 | 794.0718 | 2051 | 296 | 796 | 2079 | 292 | -2.39 | -1.01 | -1.87 | 14.432 | 1.01 |  | * | There is a smaller peak just right of the main peak. Going to calibrate like this for now.
|  |  |  | 2 | 7764 | 10729 | 12441 | 14322 | 16401 | 18621 | 14186 | 795.7624 | 13863 | 335 | 797.5 | 14015 | 329 | -0.99 | 0.11 | -0.96 | 2.417 | 0.11 |  |  | This double peak behavior only exists for 750-800 V
|  |  |  | 3 | 22708 | 30148 | 34213 | 38421 | 43062 | 48069 | 38214 | 834.9239 | 44801 | 388 | 836 | 44989 | 385 | -0.44 | -0.02 | -0.54 | 0.866 | 0.02
| LV2121 | S2B1 | 2:56PM->3:09PM | 1 | 1388 | 1930 | 2240 | 2576 | 2962 | 3375 | 2530 | 764.2112 | 2055 | 275 |  |  |  | -2.19 | -2.19 | -1.82 | 13.382 | 2.19
|  |  |  | 2 | 9869 | 13249 | 15208 | 17357 | 19674 | 22198 | 17129 | 759.7544 | 13821 | 322 |  |  |  | -1.3 | -1.3 | -1.33 | 2.33 | 1.3
|  |  |  | 3 | 29673 | 39484 | 45137 | 51305 | 56695 | 60914 | 50985 | 773.6797 | 44533 | 416 | 775.5 | 44998 | 417 | -1.05 | -0.004 | -0.63 | 0.934 | 0.004
S3T | LV2122 | S3T5 |  | 1 | 1036 | 1586 | 1891 | 2224 | 2599 | 3018 | 2238 | 790.6153 | 2109 | 237 |  |  |  | 0.43 | 0.43 | 0.63 | 11.238 | 0.43 |  | *
|  |  |  | 2 | 8004 | 11050 | 12839 | 14830 | 17015 | 19431 | 14921 | 789.8228 | 14093 | 294 |  |  |  | 0.66 | 0.66 | 0.61 | 2.086 | 0.66
|  |  | 2:16PM->2:30PM | 3 | 23801 | 31611 | 36217 | 41259 | 46792 | 52652 | 41588 | 817.1954 | 45308 | 372 | 815 | 44807 | 366 | 0.68 | -0.43 | 0.79 | 0.821 | 0.43
| LV2090 | S3T4 |  | 1 | 971 | 1535 | 1856 | 2219 | 2630 | 3091 | 2221 | 792.0631 | 2094 | 250 |  |  |  | -0.29 | -0.29 | 0.09 | 11.939 | 0.29 |  | * | LV2063 no idea why LED3 is not linear to 1 and 2?
|  |  |  | 2 | 7771 | 11151 | 13202 | 15532 | 18180 | 21178 | 15582 | 783.7688 | 14013 |  |  |  |  | 0.09 | 0.09 | 0.32 | 0 | 0.09 |  | * | LV2126 had to be moved to E21 because QDC2.021 does not work
|  |  | 2:45PM->2:55PM | 3 | 24443 | 34809 | 41267 | 48655 | 56342 | 62873 | 48860 | 787.1426 | 44889 | 405 |  |  |  | -0.25 | -0.25 | 0.42 | 0.902 | 0.25
| LV2062 | S3T3 |  | 1 | 1304 | 1989 | 2391 | 2840 | 3338 | 3902 | 2837 | 757.277 | 2104 | 252 |  |  |  | 0.19 | 0.19 | -0.11 | 11.977 | 0.19
|  |  |  | 2 | 9784 | 14041 | 16668 | 19649 | 22979 | 26498 | 19710 | 748.7443 | 14009 | 306 |  |  |  | 0.06 | 0.06 | 0.31 | 2.184 | 0.06
|  |  | 3:14PM->3:26PM | 3 | 31110 | 44114 | 51946 | 59742 | Saturates | Saturates | 60000 | 754.5437 | 45609 | 402 | 752 | 44912 | 393 | 1.34 | -0.2 | 0.43 | 0.881 | 0.2
| LV2063 | S3T2 |  | 1 | 1662 | 2435 | 2903 | 3434 | 4048 | 4723 |  | 729.884 | 2083 | 267 |  |  |  | -0.82 | -0.82 | #DIV/0! | 12.818 | 0.82 | [8 empty] | out of range | 1457 | 1773 | 2125 | 2530 | 2988 | 2120 | 798.1768 |  |  | 800 | 2120 | 224
|  |  |  | 2 | 11776 | 16821 | 19893 | 23329 | 27128 | 31153 |  | 723.2843 | 13908 | 312 |  |  |  | -0.66 | -0.66 | #DIV/0! | 2.243 | 0.66 | [8 empty] | 7291 | 10336 | 11887 | 13679 | 15402 | 17147 | 13633 | 804.4685 | 13977 | 268
|  |  | 3:47PM->3:58PM | 3 | 33089 | 45886 | 52244 | 59870 |  |  |  | 753.5108 | 46767 | 402 | 747 | 44980 | 390 | 3.78 | -0.04 | #DIV/0! | 0.86 | [9 empty] | 20459 | 25534 | 28284 | 31117 | 34083 | 37165
| LV2126 | S3T1 |  | 1 | 1204 | 1857 | 2223 | 2639 | 3106 | 3632 | 2616 | 766.95 | 2078 | 234 |  |  |  | -1.06 | -1.06 | -0.88 | 11.261 | 1.06 |  | * | QDC2.021 does not work so switched ribbon cable over to QDC1.021,
|  |  |  | 2 | 9064 | 12828 | 15112 | 17642 | 20507 | 23696 | 17546 | 763.0189 | 13920 | 294 |  |  |  | -0.57 | -0.57 | -0.55 | 2.112 | 0.57 |  |  | so the decoupler, HV, and cable remain the same just different QDC channel
|  |  | 4:31PM->4:47PM | 3 | 27493 | 38277 | 44840 | 52257 | 59410 | Saturates | 52190 | 775.0403 | 44797 | 380 | 776.5 | 45195 | 387 | -0.45 | 0.43 | -0.13 | 0.848 | 0.43
S3B | LV2136 | S3B5 |  | 1 | 1051 | 1583 | 1866 | 2185 | 2545 | 2941 |  | 792.7719 | 2102 | 327 |  |  |  | 0.1 | 0.1 | #DIV/0! | 15.557 | 0.1 |  | * | drift test after resting for the weekend for LV2136 at 800 V/LED setting2: starts at 13509 after 45 minute warm up. Now pulsing LED into LV2136 while warming up
|  |  |  | 2 | 8333 | 11436 | 13260 | 15297 | 17526 | 19982 |  | 783.4809 | 13882 | 369 |  |  |  | -0.85 | -0.85 | #DIV/0! | 2.658 | 0.85 |  | * | drift test value taken at 10:23AM 3/11 (1 hr 20 min warm up and 20 with LED)
|  |  | 8:51AM->9:11 am | 3 | 24940 | 33168 | 37938 | 43149 | 48892 | 55197 |  | 807.1484 | 44620 | 432 |  |  |  | -0.85 | -0.85 | #DIV/0! | 0.968 | 0.85 |  | * | LV2079 doesn't have much of a peak at 700 V
| LV2176 | S3B4 |  | 1 | 1092 | 1673 | 1995 | 2359 | 2774 | 3233 | 2488 | 781.2772 | 2093 | 245 | 773.7 | 2087 | 239 | -0.33 | -0.62 | 5.18 | 11.706 | 0.62 |  | * | 4/4 started warm up of LV2136 at 800V with led on set3 and still after 20 minutes the peak was only at 9128; moved motor to home and run, nothing changed
|  |  |  | 2 | 8674 | 11970 | 13881 | 16009 | 18337 | 20928 | 16443 | 775.7228 | 14336 | 292 | 773.7 | 14161 | 293 | 2.34 | 1.14 | 2.64 | 2.037 | 1.14 |  | * | trying the next PMT so moved LED to LV2176 and this one looks fine but was at 36633 and went to 44461 after 12 minutes
|  |  | 8:12:00 AM ->8:28 AM | 3 | 25778 | 34213 | 39070 | 44380 | 50157 | 56020 | 44939 | 802.6656 | 44976 | 361 |  |  |  | -0.05 | -0.05 | 1.24 | 0.803 | 0.05 |  | * | went back to LV2136 and now is ok??
| LV2079 | S3B3 |  | 1 | out of range | out of range | 1052 | 1298 | 1573 | 1826 | 1300 | 871.3862 | [5 empty] | #DIV/0! |  | 0.15 | #DIV/0! |  |  | *LV2136 started at 9241 like earlier
|  |  |  | 2 | 5717 | 8071 | 9503 | 11129 | 12959 | 15001 | 11245 | 837.7995 | 14168 | 313 |  |  |  | 1.19 | 1.19 | 1.03 | 2.209 | 1.19 |  | *700 and 750 out of range and 775 had lots of noise and low counts in the peak
|  |  | 10:05am -> 10:30am | 3 | 19544 | 27312 | 31899 | 36981 | 42552 | 48582 | 37171 | 835.3674 | 45355 | 404 | 834.5 | 45016 | 403 | 0.78 | 0.04 | 0.51 | 0.891 | 0.04
| LV2083 | S3B2 |  | 1 | 1336 | 2044 | 2451 | 2922 | 3451 | 4043 | 2945 | 753.0392 | 2092 | 282 |  |  |  | -0.38 | -0.38 | 0.78 | 13.48 | 0.38
|  |  |  | 2 | 10056 | 14459 | 17188 | 20330 | 23942 | 28120 | 20432 | 745.1849 | 14073 | 337 |  |  |  | 0.52 | 0.52 | 0.5 | 2.395 | 0.52 |  | * | LV2200 is very high gain and at 775 its still slightly out of range
|  |  | 11:03AM->11:22AM | 3 | 32454 | 46597 | 55313 | 62851 | Saturates | Saturates | 62932 | 748.0759 | 46091 | 436 | 745 | 45102 | 439 | 2.37 | 0.23 | 0.13 | 0.946 | 0.23 |  | * | LV2136 on the oscilloscope for LED3 at 800 was around 100 mV and LV2200 on LED2 at 707 is around 100 mV suggesting the PMT or decoupling channel is the problem not the DAQ
| LV2200 | S3B1 |  | 1 | 1848 | 2748 | 3302 | 3925 | 4634 | 5412 | 4004 | 715.3287 | 2150 | 308 |  |  |  | 2.33 | 2.33 | 1.97 | 14.326 | 2.33
|  |  |  | 2 | 13261 | 19139 | 22770 | 27965 | 31707 | 36887 | 27280 | 707.9941 | 14244 | 356 | 707 | 14151 | 354 | 1.71 | 1.07 | -2.51 | 2.499 | 1.07
|  |  | 11:55AM->12:05PM | 3 | 41050 | 57979 | 65187 | Saturates | Saturates | Saturates | saturates | 713.3643 | 45694 |  | 712 | 45059 | 442 | 1.52 | 0.13 | #VALUE! | 0 | 0.13
S4T | LV2146 | S4T5 |  | 1 | 1401 | 2078 | 2471 | 2913 | 3404 | 3948 | 2818 | 751.5514 | 2024 | 286 | 753 | 2070 | 286 | -3.75 | -1.45 | -3.37 | 14.13 | 1.45
|  |  |  | 2 | 9856 | 13753 | 16045 | 19624 | 21423 | 24511 | 18312 | 749.5902 | 13469 | 331 | 752.5 | 13806 | 332 | -3.94 | -1.41 | -7.16 | 2.457 | 1.41
|  |  | 3:50PM->4:05PM | 3 | 29110 | 39151 | 45011 | 51349 | 57971 | 63659 | 51016 | 774.0244 | 44467 | 393 | 776 | 44949 | 402 | -1.2 | -0.11 | -0.65 | 0.884 | 0.11 |  |  | * | LV2155 has low counts on 700 LED1
| LV2155 | S4T4 |  | 1 | 1172 | 1727 | 2011 | 2324 | 2673 | 3050 | 2280 | 781.6205 | 2046 | 296 | 784 | 2068 | 294 | -2.64 | -1.55 | -1.93 | 14.467 | 1.55 |  | * | LV2146 was a little short of yesterdays mark potentially indicating the exsistence of drift from improper warm up
|  |  |  | 2 | 8943 | 11973 | 13705 | 15594 | 17615 | 19795 | 15387 | 778.1767 | 13746 | 344 | 779.5 | 13846 | 349 | -1.85 | -1.11 | -1.35 | 2.503 | 1.11 |  | * | Going to let LV2146 warm up now and see what the 800 V and calibrated values are
|  |  | 4:23PM->4:34PM | 3 | 26520 | 34788 | 39450 | 44474 | 49787 | 54788 | 44057 | 802.1179 | 44517 | 420 | 804 | 44917 | 418 | -1.08 | -0.18 | -0.95 | 0.943 | 0.18 |  |  | 10:27->10:43 |  | was a negligible difference
| LV2169 | S4T3 |  | 1 | 966 | 1445 | 1762 | 2100 | 2480 | 2913 | 2139 | 799.6425 | 2076 | 286 |  |  |  | -1.16 | -1.16 | 1.82 | 13.776 | 1.16
|  |  |  | 2 | 7805 | 10964 | 12835 | 14955 | 17311 | 19938 | 15053 | 788.7905 | 13927 | 342 |  |  |  | -0.52 | -0.52 | 0.65 | 2.456 | 0.52
|  |  | 7:35AM -> 8:10 | 3 | 24206 | 32987 | 38039 | 43553 | 49707 | 55769 | 43514 | 805.8348 | 44812 | 455 |  |  |  | -0.42 | -0.42 | -0.09 | 1.015 | 0.42 |  |  | * | LV2169 Looks like it becomes sort of double peak st 775V and 800V, the FWHM goes from normal 421 -> 624, 640 but then it is ok at 825+
| LV2084 | S4T2 |  | 1 | 988 | 1518 | 1828 | 2172 | 2551 | 2963 | 2190 | 795.1018 | 2122 | 304 |  |  |  | 1.04 | 1.04 | 0.82 | 14.326 | 1.04
|  |  |  | 2 | 7525 | 10626 | 12477 | 14564 | 16914 | 19580 | 14675 | 793.2778 | 14066 | 360 |  |  |  | 0.47 | 0.47 | 0.76 | 2.559 | 0.47
|  |  | 8:42AM->9:02AM | 3 | 23656 | 33181 | 38716 | 45403 | 52966 | 60208 | 45528 | 797.6741 | 44877 | 444 |  |  |  | -0.27 | -0.27 | 0.27 | 0.989 | 0.27
| LV2199 | S4T1 |  | 1 | 975 | 1455 | 1760 | 2086 | 2457 | 2867 | 2141 | 801.1133 | 2160 |  |  |  |  | 2.78 | 2.78 | 2.57 | 0 | 2.78
|  |  |  | 2 | 7487 | 10501 | 12290 | 14320 | 16616 | 19159 | 14560 | 796.4997 | 14260 | 345 |  |  |  | 1.82 | 1.82 | 1.65 | 2.419 | 1.82
|  |  | 9:29AM->9:52AM | 3 | 23074 | 31390 | 36230 | 41863 | 48207 | 55033 | 42260 | 813.8474 | 45728 | 422 | 810 | 44785 | 418 | 1.59 | -0.48 | 0.94 | 0.923 | 0.48
S4B | LV2144 | S4B5 |  | 1 | 1036 | 1632 | 1959 | 2338 | 2762 | 3237 | 2381 | 790.1 | 2225 | 254 | 787.5 | 2180 | 253 | 5.62 | 3.67 | 1.81 | 11.416 | 3.67
|  |  |  | 2 | 8166 | 11611 | 13695 | 16043 | 18686 | 21677 | 16107 | 777.6 | 13980 | 312 |  |  |  | -0.14 | -0.14 | 0.4 | 2.232 | 0.14
|  |  | 7:44AM -> 8:10 AM | 3 | 24863 | 34764 | 40760 | 47614 | 55251 | 62113 | 47452 | 783.7 | 42982 | 392 | 790.5 | 44917 | 403 | -4.69 | -0.18 | -0.34 | 0.912 | 0.18
| LV2065 | S4B4 |  | 1 | 1311 | 1983 | 2374 | 2810 | 3306 | 3861 | 2821 | 757.9205 | 2106 | 260 |  |  |  | 0.28 | 0.28 | 0.39 | 12.346 | 0.28
|  |  |  | 2 | 9641 | 13735 | 16230 | 19080 | 22308 | 25914 | 19094 | 752.5198 | 13986 | 313 |  |  |  | -0.1 | -0.1 | 0.07 | 2.238 | 0.1
|  |  | 9:24AM->9:41AM | 3 | 30193 | 42667 | 50268 | 58237 | 64817 | saturates | 58244 | 756.2886 | 44501 | 395 |  |  |  | -1.12 | -1.12 | 0.01 | 0.888 | 1.12
| LV2031 | S4B3 |  | 1 | 1264 | 1926 | 2313 | 2754 | 3245 | 3799 | 2760 | 761.5933 | 2105 | 252 |  |  |  | 0.24 | 0.24 | 0.22 | 11.971 | 0.24
|  |  |  | 2 | 9838 | 14102 | 16739 | 19716 | 23115 | 26961 | 19728 | 748.6059 | 13955 | 312 |  |  |  | -0.32 | -0.32 | 0.06 | 2.236 | 0.32
|  |  | 9:53AM->10:05AM | 3 | 31270 | 44169 | 51867 | 59222 | saturates | saturates | 59217 | 755.0015 | 45605 | 406 |  |  |  | 1.33 | 1.33 | -0.01 | 0.89 | 1.33
| LV2167 | S4B2 |  | 1 | 1659 | 2341 | 2753 | 3211 | 3726 | 4283 | 3190 | 733.4791 | 2073 | 323 |  |  |  | -1.3 | -1.3 | -0.66 | 15.581 | 1.3
|  |  |  | 2 | 11560 | 15661 | 18025 | 20596 | 23343 | 26353 | 20497 | 730.4136 | 13891 | 362 |  |  |  | -0.78 | -0.78 | -0.48 | 2.606 | 0.78
|  |  | 10:20AM->10:35AM | 3 | 33660 | 44472 | 50577 | 56993 | 62331 | saturates | 56849 | 750.1731 | 44382 | 430 | 751 | 44560 | 435 | -1.39 | -0.99 | -0.25 | 0.969 | 0.99
| LV2128 | S4B1 |  | 1 | 1148 | 1782 | 2141 | 2551 | 3030 | 3552 | 2536 | 772.2288 | 2080 | 321 |  |  |  | -0.96 | -0.96 | -0.59 | 15.433 | 0.96 |  |  | I fitted without the point at 700 because I looks very off
|  |  |  | 2 | 8824 | 12589 | 14875 | 17463 | 20389 | 23696 | 17465 | 765.4371 | 13945 | 372 |  |  |  | -0.39 | -0.39 | 0.01 | 2.668 | 0.39
|  |  | 10:52AM->11:07AM | 3 | 27288 | 37952 | 44417 | 51905 | 58921 | 64992 | 51967 | 775.431 | 44586 | 447 |  |  |  | -0.93 | -0.93 | 0.12 | 1.003 | 0.93
[18 empty] | σ: | 0.9654 | % | Average Gain Accuracy: | 0.7111 | %
Note on S1: Dziubinski, Sean: % error = (experimental - actual) / actual *100

## AirVsAr
|  | 761.1 | 753.4 | 774.3
| LV1762 | LED1 | LED2 | LED3
original | Ar/air @ 720 Torr | 2097 | 13923 | 44696
| Ar @ 800 Torr | 2005 | 13725 | 45269
| Ar @ >800 Torr | 2008 | 13530 | 44174
new LED | air @ 760 Torr | 3319 | 18158 | 52332
| air @ 760 Torr | 2013 | 13530 | 44280
| LV2143 | 758.2 | 759.6 | 778.4
original | Ar/air @ 720 Torr | 2057 | 13883 | 44719
| Ar @ 800 Torr | 2180 | 14979 | 48711
new LED | air @ 760 Torr | 3421 | 19362 | 56052
old LED | air @ 760 Torr | 2238 | 15161 | 49130

## LongBoards
| PMT# | Location | Warm Up Time | LED Setting | 700 V | 750 V | 775 V | 800 V | 825 V | 850 V | Drift Test @ 800 V | Fit Voltage | Calibrated | Calibrated Resolution | Adjusted Voltage | Calibrated 2 | Calibrated 2 Resolution | Fit Accuracy (%) | Gain Accuracy (%) | Drift (%) | PMT Resolution (%) | ABS Gain
S1L | LV2194 | S1L10 | 8:45 -> 10:25AM | 1 | 1454 | 2176 | 2613 | 3103 | 3647 | 4249 | 3032 | 744.4973 | 2042 | 296 | 748 | 2091 | 295 | -2.84 | -0.43 | -2.34 | 14.496 | 0.43 |  | * 4/13/24 Now all measurements in module 0
|  |  | went back 11:35 | 2 | 9655 | 13677 | 16076 | 18788 | 21844 | 25246 | 18440 | 752.6805 | 13643 | 349 | 757 | 14047 | 347 | -2.62 | 0.33 | -1.89 | 2.558 | 0.33 |  | * 4/13/24 Now all measurements in air
|  |  | with new LED sett | 3 | 25999 | 35665 | 41771 | 48732 | 56429 | 63205 | 47820 | 786.3194 | 44007 | 437 | 790 | 45063 | 437 | -2.26 | 0.14 | -1.91 | 0.993 | 0.14 |  | * LV2194 looks double peaked at 700V on led3
| LV2140 | S1L9 | 8:45 -> 10:50 AM | 1 | 1848 | 2574 | 3002 | 3474 | 3985 | 4574 | 3388 | 718.3025 | 2034 | 288 | 722 | 2081 | 291 | -3.24 | -0.91 | -2.54 | 14.159
|  |  | with break to move | 2 | 12431 | 16772 | 19263 | 22022 | 25064 | 28358 | 21735 | 718.6267 | 13729 | 315 |  |  |  | -1.97 | -1.97 | -1.32 | 2.294
|  |  |  | 3 | 33575 | 44857 | 51309 | 57225 | 61705 | saturates | 56908 | 752.7442 | 45202 | 417 |  |  |  | 0.45 | 0.45 | -0.56 | 0.923 |  |  | * LV2140 almost saturated at 800V, bad fit; changed LED settings moving forward
| LV2075 | S1L8 | 8:45 -> 11:58 AM | 1 | 1714 | 2446 | 2888 | 3376 | 3924 | 4526 | 3304 | 726.8513 | 2033 | 261 | 730 | 2077 | 262 | -3.3 | -1.11 | -2.18 | 12.838 | 1.11
|  |  | with breaks to move | 2 | 10639 | 14467 | 16678 | 19039 | 21606 | 24371 | 18806 | 743.8178 | 13791 | 306 |  |  |  | -1.52 | -1.52 | -1.24 | 2.219 | 1.52
|  |  |  | 3 | 27020 | 35355 | 40287 | 45641 | 51391 | 56892 | 45348 | 796.5239 | 44606 | 376 | 798 | 44967 | 380 | -0.88 | -0.07 | -0.65 | 0.843 | 0.07 |  | *LV2075 double peak at 700V on led3
| LV2131 | S1L7 |  | 1 | 1881 | 2752 | 3280 | 3875 | 4547 | 5292 | 3859 | 713.3424 | 2084 | 286 |  |  |  | -0.77 | -0.77 | -0.41 | 13.724 | 0.77
|  |  |  | 2 | 12561 | 18114 | 21729 | 25629 | 30021 | 34608 | 25784 | 713.8118 | 14034 | 342 |  |  |  | 0.24 | 0.24 | 0.6 | 2.437 | 0.24
|  |  |  | 3 | 35721 | 51232 | 59265 | saturates | saturates | saturates |  | 730.1719 | 44963 | 447 |  |  |  | -0.08 | -0.08 | #VALUE! | 0.994 | 0.08
| LV2192 | S1L6 |  | 1 | 1431 | 2066 | 2432 | 2851 | 3321 | 3837 | 2788 | 751.6672 | 2036 | 256 | 756 | 2091 |  | -3.14 | -0.43 | -2.26 | 12.574 | 0.43 |  | * LV2187 double peak at 700V on led3
|  |  |  | 2 | 9687 | 13304 | 15430 | 17798 | 20382 | 23245 | 17541 | 757.6356 | 13739 | 298 |  |  |  | -1.9 | -1.9 | -1.47 | 2.169 | 1.9
|  |  |  | 3 | 25579 | 33983 | 38901 | 44289 | 50103 | 56207 | 43875 | 802.6428 | 44499 | 368 | 806 | 45287 | 373 | -1.13 | 0.63 | -0.94 | 0.827 | 0.63
| LV2165 | S1L5 | 2:50PM-> 3:55 PM | 1 | 1607 | 2380 | 2843 | 3365 | 3949 | 4586 | 3336 | 732.0843 | 2055 | 304 | 736 | 2122 | 305 | -2.19 | 1.04 | -0.87 | 14.793 | 1.04
|  |  |  | 2 | 10481 | 14992 | 17784 | 21031 | 24787 | 28893 | 20854 | 739.7103 | 13857 | 347 |  |  |  | -1.03 | -1.03 | -0.85 | 2.504 | 1.03
|  |  |  | 3 | 29776 | 42650 | 50956 | 59845 | saturate | saturate | 59675 | 756.8103 | 44716 | 456 |  |  |  | -0.64 | -0.64 |  | 1.02 | 0.64
| LV2087 | S1L4 | 700, 725, 750, 775, 800 | 1 | 2215 | 2706 | 3271 | 3916 | 4644 |  |  | 693.4983 | 2109 | 317 |  |  |  | 0.43 | 0.43 |  | 15.031 | 0.43 |  | 2:50 ->4:10 PM | 1 | too low | 1226 | 1492 | 1735 | 1992 | 2266 | 1721 | 833.7596 | 2065 | 264
|  |  |  | 2 | 15245 | 17965 | 20892 | 24074 | 27351 |  |  | 697.5728 | 15190 | 342 | 688 | 14202 | 337 | 7.83 | 1.42 |  | 2.251 | 1.42 |  | older settings LED | 2 | 6089 | 8132 | 9302 | 10571 | 11921 | 13380 | 10529 | 858.4943
|  |  |  | 3 | 38613 | 44708 | 51421 | 58497 | 64417 |  |  | 725.9655 | 45100 | 393 |  |  |  | 0.22 | 0.22 | -0.17 | #DIV/0! | 0.22 |  | still too low overall | 3 | 16768 | 21829 | 24619 | 27613 | 30730 | 34056 | 27567 | 921.8182
| LV2064 | S1L3 |  | 1 | 1306 | 1975 | 2363 | 2805 | 3288 | 3835 | 2743 | 757.5547 | 2047 | 262 | 761 | 2103 | 263 | -2.59 | 0.14 | -2.26 | 12.799 | 0.14
|  |  |  | 2 | 9065 | 12958 | 15326 | 18029 | 21117 | 24637 | 17855 | 760.7756 | 13792 | 313 |  |  |  | -1.51 | -1.51 | -0.97 | 2.269 | 1.51
|  |  |  | 3 | 26026 | 37166 | 43727 | 51708 | 58746 | 64796 | 51537 | 776.773 | 44151 | 428 | 780 | 45158 |  | -1.92 | 0.35 | -0.33 | 0.969 | 0.35 |  | * LV2064 double peak at 750V
| LV2117 | S1L2 |  | 1 | 1998 | 2843 | 3354 | 3921 | 4551 | 5257 | 3890 | 706.0927 | 2067 | 295 | 708 | 2098 | 294 | -1.6 | -0.1 | -0.8 | 14.272 | 0.1
|  |  |  | 2 | 12490 | 16940 | 19500 | 22290 | 25306 | 28573 | 22265 | 717.4635 | 13923 | 386 |  |  |  | -0.55 | -0.55 | -0.11 | 2.772 | 0.55
|  |  |  | 3 | 32188 | 42638 | 48614 | 55048 | 60537 | saturates | 55175 | 758.8831 | 44778 | 407 |  |  |  | -0.5 | -0.5 | 0.23 | 0.909 | 0.5
| LV2098 | S1L1 | finished points ~5:20PM | 1 | 1737 | 2476 | 2927 | 3433 | 3991 | 4623 | 3385 | 725.1813 | 2049 | 268 | 728 | 2098 | 266 | -2.49 | -0.1 | -1.42 | 13.08 | 0.1 |  |  | 9:36
|  |  | finished calib 6:10 | 2 | 11418 | 15681 | 18182 | 20945 | 23990 | 27278 | 20768 | 730.4505 | 13774 | 311 | 733 | 14011 | 309 | -1.64 | 0.08 | -0.85 | 2.258 | 0.08
|  |  |  | 3 | 29191 | 38731 | 44127 | 49982 | 56182 | 61719 | 49864 | 777.4985 | 44580 | 371 | 779 | 44984 | 367 | -0.94 | -0.04 | -0.24 | 0.832 | 0.04
S1R | LV2100 | S1R10 | 11:40AM->12:58PM | 1 | 1666 | 2471 | 2962 | 3506 | 4112 | 4789 | 3526 | 727.4912 | 2068 | 309 | 728.5 | 2083 | 304 | -1.55 | -0.82 | 0.57 | 14.942 | 0.82
|  |  |  | 2 | 11140 | 15479 | 17988 | 20784 | 23802 | 27070 | 20744 | 733.8162 | 13868 | 338 |  |  |  | -0.95 | -0.95 | -0.19 | 2.437 | 0.95
|  |  |  | 3 | 28947 | 38883 | 44662 | 50967 | 57572 | 63074 | 50900 | 775.1874 | 44465 | 418 | 777 | 44997 | 418 | -1.2 | -0.01 | -0.13 | 0.94 | 0.01
| LV1785 | S1R9 | 11:40AM->1:40PM | 1 | 981 | 1466 | 1788 | 2140 | 2547 | 3002 | 2147 | 797.5049 | 2108 | 270 |  |  |  | 0.38 | 0.38 | 0.33 | 12.808 | 0.38
|  |  |  | 2 | 7049 | 10078 | 11916 | 13969 | 16278 | 18888 | 14083 | 800.3683 | 14004 | 320 |  |  |  | 0.03 | 0.03 | 0.81 | 2.285 | 0.03
|  |  |  | 3 | 20049 | 28130 | 32962 | 38281 | 44049 | 50863 | 38394 | 828.4847 | 44924 | 416 |  |  |  | -0.17 | -0.17 | 0.29 | 0.926 | 0.17
| LV2091 | S1R8 | 11:40->2:04PM | 1 | 1208 | 1857 | 2219 | 2636 | 3104 | 3625 | 2579 | 766.9121 | 2096 | 266 |  |  |  | -0.19 | -0.19 | -2.21 | 12.691 | 0.19
|  |  |  | 2 | 8425 | 12020 | 14222 | 16724 | 19581 | 22799 | 16595 | 772.7619 | 14003 | 309 |  |  |  | 0.02 | 0.02 | -0.78 | 2.207 | 0.02
|  |  |  | 3 | 23836 | 33885 | 39879 | 46767 | 54484 | 60885 | 46825 | 793.8023 | 44829 | 411 |  |  |  | -0.38 | -0.38 | 0.12 | 0.917 | 0.38
| LV2095 | S1R7 | 11:40AM->2:37PM | 1 | 1282 | 1966 | 2357 | 2801 | 3303 | 3865 | 2722 | 758.9918 | 2103 | 290 |  |  |  | 0.14 | 0.14 | -2.9 | 13.79 | 0.14
|  |  |  | 2 | 9121 | 13165 | 15659 | 18534 | 21859 | 25657 | 18266 | 758.8266 | 13966 | 343 |  |  |  | -0.24 | -0.24 | -1.47 | 2.456 | 0.24
|  |  |  | 3 | 26842 | 38297 | 45360 | 53423 | 60386 | saturates | 53325 | 773.1313 | 44694 | 440 | 775 | 44991 | 847 | -0.68 | -0.02 | -0.18 | 0.984 | 0.02
| LV2187 | S1R6 | 11:40AM->3:02PM | 1 | 1202 | 1813 | 2152 | 2539 | 2965 | 3441 | 1951 | 771.0658 | 2092 | 265 |  |  |  | -0.38 | -0.38 | -30.14 | 12.667 | 0.38
|  |  |  | 2 | 8621 | 11977 | 13979 | 16225 | 18729 | 21512 | 12709 | 775.1053 | 13923 | 310 |  |  |  | -0.55 | -0.55 | -27.67 | 2.227 | 0.55 | checked S1R10 again after this measurement and its still the same so its not the LED must be PMT or QDC related
|  |  |  | 3 | 23888 | 32148 | 37039 | 42714 | 49006 | 55805 | 34321 | 809.6504 | 44910 | 389 |  |  |  | -0.2 | -0.2 | -24.45 | 0.866 | 0.2 | double peak at 800 drift
| LV2190 | S1R5 | 7:34AM->10:43AM | 1 | 1321 | 1939 | 2291 | 2685 | 3121 | 3613 | 2626 | 761.5782 | 2089 | 263 |  |  |  | -0.53 | -0.53 | -2.25 | 12.59 | 0.53
|  |  |  | 2 | 8993 | 12327 | 14271 | 16432 | 18788 | 21402 | 16313 | 771.4103 | 13937 | 302 |  |  |  | -0.45 | -0.45 | -0.73 | 2.167 | 0.45
|  |  |  | 3 | 23852 | 31460 | 35963 | 40884 | 46241 | 52080 | 40796 | 819.5376 | 44908 | 374 |  |  |  | -0.2 | -0.2 | -0.22 | 0.833 | 0.2
| LV2186 | S1R4 | 600, 625, 650, 675, 700 | 1 |  | 1457 | 1793 | 2172 | 2615 | 3116 |  | 645.336 | 2146 | 295 | 644 | 2124 | 293 | 2.14 | 1.13 |  | 13.747 | 1.13 |  | 3099 | 4379 | 5130 | 5956 |  |  | 5885 | 652.013
|  |  |  | 2 |  | 10385 | 12439 | 14734 | 17273 | 20036 |  | 641.6374 | 14025 | 329 |  |  |  | 0.18 | 0.18 |  | 2.346 | 0.18 |  | 18187 | 24219 | 27603 | 31283 | 35269 | 39594 | 31167 | 659.7729
|  |  |  | 3 |  | 27706 | 32634 | 38004 | 43858 | 50151 |  | 679.6983 | 45105 | 397 |  |  |  | 0.23 | 0.23 |  | 0.88 | 0.23 |  | 44543 | 57775 | 63524 |  |  |  |  | 702.0771 | 45210 | 397
| LV2077 | S1R3 | 7:34 AM -> 10:01 AM | 1 | 1293 | 1941 | 2310 | 2729 | 3197 | 3703 | 2752 | 761.1108 | 2107 | 261 |  |  |  | 0.33 | 0.33 | 0.84 | 12.387 | 0.33
|  |  |  | 2 | 8838 | 12247 | 14191 | 16368 | 18733 | 21293 | 16498 | 772.3067 | 13925 | 304 |  |  |  | -0.54 | -0.54 | 0.79 | 2.183 | 0.54
|  |  |  | 3 | 23551 | 31475 | 35773 | 40504 | 45723 | 51278 | 40724 | 821.5786 | 44815 | 382 |  |  |  | -0.41 | -0.41 | 0.54 | 0.852 | 0.41
| LV1760 | S1R2 | 7:34AM->9:33AM | 1 | 1771 | 2569 | 3060 | 3619 | 4242 | 4944 | 3643 | 721.9524 | 2088 | 271 |  |  |  | -0.57 | -0.57 | 0.66 | 12.979 | 0.57
|  |  |  | 2 | 11686 | 16207 | 18861 | 21792 | 24950 | 28366 | 21949 | 726.8767 | 13998 | 316 |  |  |  | -0.01 | -0.01 | 0.72 | 2.257 | 0.01
|  |  |  | 3 | 30518 | 41011 | 46958 | 53446 | 59284 | 64292 | 53635 | 765.9099 | 44597 | 391 | 766.9437 | 44881 | 395 | -0.9 | -0.27 | 0.35 | 0.877 | 0.27
| LV2123 | S1R1 | 7:34AM->8:58AM | 1 | 1709 | 2485 | 2969 | 3507 | 4114 | 4768 | 3633 | 726.2097 | 2099 | 309 |  |  |  | -0.05 | -0.05 | 3.47 | 14.721 | 0.05
|  |  |  | 2 | 11434 | 15653 | 18076 | 20728 | 23534 | 26498 | 20953 | 731.5064 | 13976 | 337 |  |  |  | -0.17 | -0.17 | 1.07 | 2.411 | 0.17
|  |  |  | 3 | 29079 | 38704 | 44194 | 50150 | 56605 | 62572 | 50258 | 778.3853 | 44786 | 411 |  |  |  | -0.48 | -0.478 | 0.21 | 0.918 | 0.478
S2L | LV1780 | S2L10 | 9:35:00 AM->10:26AM | 1 | 1518 | 2226 | 2655 | 3148 | 3693 | 4300 |  | 741.6712 | [5 empty] | 1.32 | 1.32 | 1.82 | 16.212 | 1.32
|  | 2024-04-14 | went back 13:15 | 2 | 10664 | 14876 | 17380 | 20157 | 23270 | 26709 |  | 740.4458 | [5 empty] | 4.64 | 0.48 | 2.28 | 2.398 | 0.48
|  |  |  | 3 | 29003 | 39002 | 44817 | 51216 | 57844 | 63561 |  | 774.7556 | [5 empty] | 0.39 | 0.39 | #VALUE! | 0.996 | 0.39 | 650, 675, 700, 725 |  | 39572 | 48720 | 59225 | saturates
| LV1789 | S2L9 | 9:35:00 AM->11:27AM | 1 | 1441 | 2191 | 2636 | 3152 | 3670 | 4412 | 3175 | 745.9069 | 2112 | 272 |  |  |  | 0.57 | 0.57 | 0.72 | 12.879 | 0.57 |  |  | 2995 | 4432 | 5310 | 6313 | 7428 | 8677 | 6430 | 661.6998 | 2128 | 345
|  |  | went back 13:18 | 2 | 9943 | 14391 | 17146 | 20298 | 23875 | 27923 | 20644 | 745.7998 | 13996 | 321 |  |  |  | -0.03 | -0.03 | 1.68 | 2.294 | 0.03 |  |  | 20604 | 30101 | 36030 | 43020 | 50593 | 57818 | 44025 | 659.9261 | 14681 | 352 | 654 | 14067 | 348
|  |  |  | 3 | 28774 | 41021 | 48834 | 57086 | 63693 | saturated | 57071 | 763.2526 | 44933 | 434 |  |  |  | -0.15 | -0.15 | -0.03 | 0.966 | 0.15 | exclude 825 V point from the fitting - better fitting |  | 59225 | 71350 | 85132 | saturates |  |  |  | 665.2688 | 45178 | 450
| LV2179 | S2L8 | 9:35:00 AM->11:55AM | 1 | 1248 | 1887 | 2252 | 2660 | 3126 | 3644 | 2680 | 765.1332 | 2112 | 264 |  |  |  | 0.57 | 0.57 | 0.75 | 12.5 | 0.57
|  |  | went back 13:22 | 2 | 8975 | 12512 | 14615 | 16944 | 19542 | 22440 | 17042 | 768.1515 | 13980 | 312 |  |  |  | -0.14 | -0.14 | 0.58 | 2.232 | 0.14
|  |  |  | 3 | 24850 | 33760 | 38714 | 44363 | 50804 | 57032 | 44325 | 802.3881 | 44804 | 395 | 803 | 45098 | 388 | -0.44 | 0.22 | -0.09 | 0.882 | 0.22
| LV2173 | S2L7 | 9:35:00 AM->12:25PM | 1 | 1163 | 1746 | 2070 | 2432 | 2835 | 3274 | 2436 | 776.7442 | 2093 | 287 |  |  |  | -0.33 | -0.33 | 0.16 | 13.712 | 0.33
|  |  | went back 13:25 | 2 | 8556 | 11715 | 13550 | 15600 | 17855 | 20287 | 15656 | 780.7509 | 14028 | 325 |  |  |  | 0.2 | 0.2 | 0.36 | 2.317 | 0.2
|  |  |  | 3 | 23443 | 31154 | 35393 | 40010 | 45247 | 50956 | 40113 | 823.2192 | 44896 | 394 |  |  |  | -0.23 | -0.23 | 0.26 | 0.878
| LV2166 | S2L6 | 9:35:00 AM->12:50PM | 1 | 1238 | 1857 | 2210 | 2610 | 3056 | 3545 | 2595 | 767.4519 | 2092 | 260 |  |  |  | -0.38 | -0.38 | -0.58 | 12.428 | 0.38
|  |  | went back 13:28 | 2 | 9170 | 12653 | 14698 | 16990 | 19495 | 22216 | 16902 | 766.294 | 13945 | 304 |  |  |  | -0.39 | -0.39 | -0.52 | 2.18 | 0.39
|  |  |  | 3 | 24503 | 32488 | 37087 | 42052 | 47442 | 53193 | 42036 | 813.548 | 44924 | 367 |  |  |  | -0.17 | -0.17 | -0.04 | 0.817 | 0.17
| LV2157 | S2L5 | 14:00 PM->16:30PM | 1 | 1495 | 2153 | 2552 | 2995 | 3488 | 4036 | 3006 | 745.7176 | 2090 | 744 |  |  |  | -0.48 | -0.48 | 0.37 | 35.598 | 0.48
|  |  |  | 2 | 10077 | 13843 | 16038 | 18483 | 21211 | 24217 | 18500 | 752.2475 | 14020 | 303 |  |  |  | 0.14 | 0.14 | 0.09 | 2.161 | 0.14
|  |  |  | 3 | 26292 | 35100 | 40127 | 45605 | 51567 | 57665 | 45650 | 797.3487 | 45010 | 376 |  |  |  | 0.02 | 0.02 | 0.1 | 0.835 | 0.02
| LV2082 | S2L4 | 14:00 PM->16:05PM | 1 | 1390 | 2028 | 2402 | 2828 | 3264 | 3763 | 2818 | 755.2086 | 2104 | 288 |  |  |  | 0.19 | 0.19 | -0.35 | 13.688 | 0.19
|  |  | went back 17:10 | 2 | 9434 | 12773 | 14000 | 16753 | 18979 | 21333 | 16804 | 770.6225 | 14350 | 327 | 767 | 14060 | 322 | 2.44 | 0.43 | 0.3 | 2.279 | 0.43
|  |  |  | 3 | 24085 | 31660 | 35900 | 40490 | 45423 | 50760 | 40566 | 822.3552 | 44908 | 396 |  |  |  | -0.2 | -0.2 | 0.19 | 0.882 | 0.2
| LV2072 | S2L3 | 14:00 PM->15:40PM | 1 | 951 | 1502 | 1814 | 2158 | 2382 | 2990 | 2191 | 800.2751 | 2160 | 261 | 798 | 2120 | 264 | 2.78 | 0.94 | 1.51 | 12.083
|  |  | went back 17:05 | 2 | 6906 | 9807 | 11572 | 13527 | 15737 | 18214 | 13707 | 806.208 | 14057 | 319 |  |  |  | 0.41 | 0.41 | 1.31 | 2.269 | 0.41
|  |  |  | 3 | 19566 | 27551 | 32368 | 37610 | 43376 | 49993 | 37831 | 830.5709 | 44760 | 396 | 830 | 44900 | 402 | -0.54 | -0.22 | 0.58 | 0.885 | 0.22
| LV2058 | S2L2 | 14:00 PM->15:20PM | 1 | 1307 | 2002 | 2410 | 2864 | 3370 | 3975 | 2907 | 756.9207 | 2119 | 272 |  |  |  | 0.9 | 0.9 | 1.48 | 12.836 | 0.9
|  |  | went back 17:00 | 2 | 9598 | 13828 | 16380 | 19353 | 22705 | 26570 | 19520 | 751.7443 | 13959 | 323 |  |  |  | -0.29 | -0.29 | 0.86 | 2.314 | 0.29
|  |  |  | 3 | 28041 | 39772 | 47110 | 54950 | 61421 | saturated | 55182 | 768.0044 | 44998 | 433 |  |  |  | 0 | -0.0044 | 0.42 | 0.962 | 0.0044
| LV2161 | S2L1 | 14:00 PM->14:50PM | 1 | 1085 | 1681 | 2054 | 2390 | 2843 | 3336 | 2322 | 779.2186 | 2136 | 299 |  |  |  | 1.69 | 1.69 | -2.93 | 13.998 | 1.69
|  |  | went back 16:55 | 2 | 7521 | 10680 | 12675 | 14885 | 17268 | 19957 | 14924 | 791.2514 | 14321 | 353 | 789 | 14059 | 347 | 2.24 | 0.42 | 0.26 | 2.465 | 0.42
|  |  |  | 3 | 19937 | 27403 | 31976 | 37200 | 43070 | 50060 | 38860 | 831.8579 | 45260 | 460 | 831 | 44941 | 462 | 0.57 | -0.13 | 4.27 | 1.016 | 0.13
S4L | LV2093 | S4L10 | 7:35:00 AM ->8:50 am | 1 | 1258 | 1953 | 2364 | 2820 | 3334 | 3903 | 2927 | 758.2368 | 2086 | 298 |  |  |  | -0.67 | -0.67 | 3.66 | 14.286 | 0.67
|  |  | 2024-04-19 | 2 | 9140 | 13073 | 15454 | 18183 | 21305 | 24793 | 18421 | 759.4713 | 13928 | 343 |  |  |  | -0.52 | -0.52 | 1.29 | 2.463 | 0.52
|  |  |  | 3 | 26354 | 37184 | 44176 | 52275 | 60529 | saturated | 52439 | 778.0033 | 44973 | 452 |  |  |  | -0.06 | -0.06 | 0.31 | 1.005 | 0.06
| LV2052 | S4L9 |  | 1 | 1310 | 1988 | 2379 | 2874 | 3333 | 3895 | 2770 | 757.4295 | 2100 | 265 |  |  |  | 0 | 0 | -3.75 | 12.619 | 0
|  |  |  | 2 | 9430 | 13445 | 15899 | 18706 | 21859 | 25450 | 18441 | 755.7587 | 13944 | 314 |  |  |  | -0.4 | -0.4 | -1.44 | 2.252 | 0.4
|  |  |  | 3 | 27266 | 38431 | 45465 | 53322 | 60121 | saturates | 53203 | 774.1898 | 45293 | 424 |  |  |  | 0.65 | 0.6469 | -0.22 | 0.936 | 0.6469
| LV2078 | S4L8 | 7:35AM->10:02AM | 1 | 1118 | 1749 | 2092 | 2484 | 2926 | 3436 | 2410 | 775.289 | 2073 | 262 |  |  |  | -1.3 | -1.3 | -3.07 | 12.639 | 1.3
|  |  |  | 2 | 8001 | 11419 | 13474 | 15769 | 18334 | 21182 | 15661 | 780.9687 | 14066 | 309 |  |  |  | 0.47 | 0.47 | -0.69 | 2.197 | 0.47
|  |  |  | 3 | 22807 | 31857 | 37023 | 42445 | 48767 | 55249 | 42248 | 810.0257 | 44980 | 394 |  |  |  | -0.04 | -0.04 | -0.47 | 0.876 | 0.04
| LV2174 | S4L7 | 7:35AM->10:19AM | 1 | 1390 | 1927 | 2234 | 2578 | 2951 | 3361 | 2651 | 764.2855 | 2100 | 295 |  |  |  | 0 | 0 | 2.75 | 14.048 | 0
|  |  |  | 2 | 9784 | 13080 | 14940 | 16999 | 19203 | 21587 | 17538 | 762.3572 | 13905 | 328 |  |  |  | -0.68 | -0.68 | 3.07 | 2.359 | 0.68
|  |  |  | 3 | 26903 | 35210 | 39848 | 44856 | 50149 | 55250 | 46417 | 800.9737 | 44740 | 418 |  |  |  | -0.58 | -0.58 | 3.36 | 0.934 | 0.58
| LV2164 | S4L6 | 7:35AM->10:40AM | 1 | 1322 | 1984 | 2365 | 2797 | 3280 | 3820 | 2751 | 758.0374 | 2094 | 263 |  |  |  | -0.29 | -0.29 | -1.67 | 12.56 | 0.29
|  |  |  | 2 | 9419 | 13094 | 15282 | 17732 | 20454 | 23502 | 17627 | 760.549 | 13920 | 305 |  |  |  | -0.57 | -0.57 | -0.6 | 2.191 | 0.57
|  |  |  | 3 | 25738 | 34568 | 39965 | 45924 | 52663 | 59178 | 45880 | 796.3092 | 44774 | 389 |  |  |  | -0.5 | -0.5 | -0.1 | 0.869 | 0.5
| LV2030 | S4L5 | 11:40AM->2:30PM | 1 | 1143 | 1727 | 2065 | 2448 | 2881 | 3367 | 2402 | 777.173 | 2090 | 268 |  |  |  | -0.48 | -0.48 | -1.92 | 12.823 | 0.48
|  |  |  | 2 | 8165 | 11594 | 13607 | 15973 | 18585 | 21518 | 15794 | 779.3176 | 13963 | 305 |  |  |  | -0.26 | -0.26 | -1.13 | 2.184 | 0.26
|  |  |  | 3 | 22836 | 31699 | 37159 | 43296 | 50200 | 57290 | 43194 | 806.797 | 44947 | 404 |  |  |  | -0.12 | -0.12 | -0.24 | 0.899 | 0.12
| LV2163 | S4L4 | 11:40AM->2:11PM | 1 | 1739 | 2515 | 2982 | 3510 | 4095 | 4764 | 3479 | 724.7216 | 2095 | 296 |  |  |  | -0.24 | -0.24 | -0.89 | 14.129 | 0.24
|  |  |  | 2 | 11293 | 15521 | 18002 | 20777 | 23813 | 27186 | 20717 | 733.1891 | 13977 | 320 |  |  |  | -0.16 | -0.16 | -0.29 | 2.289 | 0.16
|  |  |  | 3 | 28816 | 38126 | 43570 | 49563 | 55859 | 61394 | 49535 | 780.8492 | 44855 | 382 |  |  |  | -0.32 | -0.32 | -0.06 | 0.852 | 0.32
| LV2066 | S4L3 | 11:40AM->1:43PM | 1 | 1285 | 1970 | 2365 | 2814 | 3316 | 3876 | 2700 | 758.6747 | 2093 | 260 |  |  |  | -0.33 | -0.33 | -4.22 | 12.422 | 0.33
|  |  |  | 2 | 8937 | 12726 | 15018 | 17640 | 20626 | 23807 | 17290 | 763.9529 | 13956 | 309 |  |  |  | -0.32 | -0.32 | -2.02 | 2.214 | 0.32
|  |  |  | 3 | 25408 | 36060 | 42338 | 49738 | 57007 | 63085 | 49182 | 784.1957 | 44645 | 415 | 785.7 | 45089 | 421 | -0.8 | -0.8 | -1.13 | 0.93 | 0.8
| LV2067 | S4L2 | 11:40AM->1:14PM | 1 | 1292 | 1989 | 2392 | 2845 | 3356 | 3924 | 2856 | 757.2556 | 2100 | 270 |  |  |  | 0 | 0 | 0.39 | 12.857 | 0
|  |  |  | 2 | 9093 | 12782 | 14946 | 17353 | 19968 | 22870 | 17397 | 764.3908 | 14506 | 314 | 762 | 14178 | 311 | 3.49 | 1.26 | 0.25 | 2.165 | 1.26
|  |  |  | 3 | 24767 | 33468 | 38544 | 44253 | 50448 | 56834 | 44430 | 802.5388 | 45143 | 389 |  |  |  | 0.32 | 0.32 | 0.4 | 0.862 | 0.32
| LV2094 | S4L1 | 11:40AM->12:47PM | 1 | 1496 | 2140 | 2525 | 2948 | 3436 | 3967 | 2954 | 746.9435 | 2094 | 300 |  |  |  | -0.29 | -0.29 | 0.2 | 14.327 | 0.29
|  |  |  | 2 | 10006 | 13613 | 16658 | 17829 | 20109 | 22575 | 17788 | 748.6781 | 13484 | 503 | 753 | 13810 | 511 | -3.83 | -1.38 | -0.23 | 3.73 | 1.38
|  |  |  | 3 | 29945 | 34283 | 39025 | 44188 | 49763 | 55730 | 44208 | 806.8909 | 45578 | 415 | 804 | 45012 | 417 | 1.27 | 0.03 | 0.05 | 0.911 | 0.03
S3L | LV2106 | S3L10 | 12:00 pm ->1:03 pm | 1 | 1613 | 2436 | 2941 | 3520 | 4173 | 4883 | 3523 | 730.4479 | 2075 | 303 |  |  |  | -1.2 | -1.2 | 0.09 | 14.602 | 1.2
|  |  | went back 3:35 pm | 2 | 11176 | 16003 | 18933 | 22286 | 26067 | 30342 | 22217 | 730.327 | 13879 | 344 |  |  |  | -0.87 | -0.87 | -0.31 | 2.479 | 0.87
|  |  |  | 3 | 30702 | 42984 | 50718 | 59128 | sat | sat | 58874 | 756.6546 | 44866 | 426 |  |  |  | -0.3 | -0.3 | -0.43 | 0.949 | 0.3
| LV2145 | S3L9 | went back 3:42 | 1 | 1455 | 2156 | 2579 | 3058 | 3589 | 4187 | 3081 | 745.3723 | 2085 | 271 |  |  |  | -0.72 | -0.72 | 0.75 | 12.998 | 0.72
|  |  |  | 2 | 9823 | 13813 | 16184 | 18831 | 21806 | 25107 | 18874 | 751.7509 | 13955 | 311 |  |  |  | -0.32 | -0.32 | 0.23 | 2.229 | 0.32
|  |  |  | 3 | 26696 | 36725 | 42586 | 49529 | 56667 | 62654 | 49685 | 783.2455 | 44861 | 394 |  |  |  | -0.31 | -0.31 | 0.31 | 0.878 | 0.31
| LV2069 | S3L8 | went back 3:48 | 1 | 1365 | 2070 | 2477 | 2953 | 3471 | 4066 | 2951 | 751.3372 | 2084 | 263 |  |  |  | -0.77 | -0.77 | -0.07 | 12.62 | 0.77
|  |  |  | 2 | 9557 | 13377 | 15672 | 18138 | 20849 | 23800 | 18147 | 755.9846 | 13752 | 304 |  |  |  | -1.8 | -1.8 | 0.05 | 2.211 | 1.8
|  |  |  | 3 | 25935 | 35136 | 40312 | 46249 | 52624 | 58354 | 46286 | 793.9178 | 44498 | 382 | 796 | 45028 | 386 | -1.13 | 0.06 | 0.08 | 0.858 | 0.06 |  | * double peak at 750V
| LV2171 | S3L7 |  | 1 | 1516 | 2196 | 2599 | 3052 | 3558 | 4125 | 3061 | 742.6778 | 2086 | 291 |  |  |  | -0.67 | -0.67 | 0.29 | 13.95 | 0.67
|  |  |  | 2 | 10387 | 14223 | 16458 | 18930 | 21640 | 24628 | 18945 | 746.3121 | 13824 | 324 |  |  |  | -1.27 | -1.27 | 0.08 | 2.344 | 1.27
|  |  |  | 3 | 27450 | 36064 | 41152 | 46675 | 52673 | 58336 | 46757 | 792.5195 | 44711 | 388 |  |  |  | -0.65 | -0.65 | 0.18 | 0.868 | 0.65
| LV2147 | S3L6 | went back 4:00 | 1 | 1296 | 1922 | 2275 | 2681 | 3124 | 3627 | 2738 | 762.6487 | 2090 | 264 |  |  |  | -0.48 | -0.48 | 2.08 | 12.632 | 0.48
|  |  | finished taking points | 2 | 9291 | 12755 | 14771 | 17006 | 19469 | 22152 | 17188 | 764.7604 | 13874 | 303 |  |  |  | -0.91 | -0.91 | 1.06 | 2.184 | 0.91
|  |  | 15:30 | 3 | 25069 | 33230 | 38048 | 43330 | 49072 | 55245 | 43523 | 806.7418 | 44739 | 379 |  |  |  | -0.58 | -0.58 | 0.44 | 0.847 | 0.58
| LV2159 | S3L5 |  | 1 | 1057 | 1556 | 1849 | 2178 | 2545 | 2951 |  | 794.2196 | 2102 | 259 |  |  |  | 0.1 | 0.1 |  | 12.322 | 0.1 |  | 1125 | 1627 | 1906 | 2216 | 2566 | 2945 | 2111 | 790.2253 | 2081 | 269
|  |  |  | 2 | 7951 | 10899 | 12618 | 14533 | 16599 | 18896 |  | 793.4053 | 13960 | 303 |  |  |  | -0.29 | -0.29 |  | 2.17 | 0.29 |  | 7604 | 10306 | 11862 | 13548 | 15374 | 17321 | 13329 | 806.1797 | 13964 | 301
|  |  |  | 3 | 21834 | 28816 | 32776 | 37152 | 41903 | 47026 |  | 832.7839 | 43220 | 354 | 840 | 44738 | 372 | -4.12 | -0.59 |  | 0.819 | 0.59 |  | 20096 | 26300 | 29661 | 33034 | 36789 | 40883 | 32794 | 874.0186
| LV2162 | S3L4 |  | 1 | 1631 | 2378 | 2875 | 3381 | 3932 | 4557 | 3248 | 731.3063 | 2162 | 294 |  |  |  | 2.87 | 2.87 | -4.09 | 13.599 | 2.87
|  |  |  | 2 | 10988 | 15120 | 17545 | 20217 | 23179 | 26414 | 20226 | 736.5985 | 13930 | 320 |  |  |  | -0.5 | -0.5 | 0.04 | 2.297 | 0.5
|  |  |  | 3 | 28782 | 38270 | 43725 | 49684 | 55838 | 61108 | 49761 | 779.2567 | 44705 | 375 |  |  |  | -0.66 | -0.66 | 0.15 | 0.839 | 0.66
| LV2054 | S3L3 |  | 1 | 1231 | 1895 | 2273 | 2707 | 3198 | 3757 | 2685 | 763.7562 | 2096 | 261 |  |  |  | -0.19 | -0.19 | -0.82 | 12.452 | 0.19
|  |  |  | 2 | 8984 | 12850 | 15182 | 17813 | 20772 | 24079 | 17749 | 762.2104 | 13966 | 311 |  |  |  | -0.24 | -0.24 | -0.36 | 2.227 | 0.24
|  |  |  | 3 | 25797 | 35918 | 41807 | 48549 | 55413 | 61103 | 48393 | 786.9017 | 44892 | 393 |  |  |  | -0.24 | -0.24 | -0.32 | 0.875 | 0.24
| LV2047 | S3L2 |  | 1 | 1327 | 2006 | 2404 | 2853 | 3357 | 3918 | 2758 | 755.4356 | 2089 | 272 |  |  |  | -0.53 | -0.53 | -3.44 | 13.021 | 0.53
|  |  |  | 2 | 9230 | 13224 | 15677 | 18425 | 21653 | 25302 | 18169 | 758.0484 | 13991 | 318 |  |  |  | -0.06 | -0.06 | -1.41 | 2.273 | 0.06
|  |  |  | 3 | 26353 | 37315 | 44373 | 52839 | 59975 | saturates | 52713 | 777.9184 | 45611 | 426 | 776 | 45009 | 423 | 1.34 | 0.02 | -0.24 | 0.934 | 0.02
| LV2139 | S3L1 | 7:30:00 AM -> 8:33 | 1 | 1805 | 2676 | 3221 | 3816 | 4492 | 5231 | 4102 | 717.4544 | 2075 | 295 |  |  |  | -1.2 | -1.2 | 6.97 | 14.217 | 1.2
|  |  |  | 2 | 12178 | 17317 | 20503 | 24120 | 28149 | 32435 | 24901 | 718.9496 | 13880 | 378 |  |  |  | -0.86 | -0.86 | 3.14 | 2.723 | 0.86
|  |  |  | 3 | 33312 | 46978 | 55367 | 63363 | saturates | saturates | 63738 | 743.3354 | 44785 | 419 |  |  |  | -0.48 | -0.48 | 0.59 | 0.936 | 0.48
S3R | LV2120 | S3R10 | 12:12PM->1:16PM | 1 | 1138 | 1753 | 2103 | 2498 | 2943 | 3427 | 2430 | 774.5541 | 2095 | 324 |  |  |  | -0.24 | -0.24 | -2.8 | 15.465 | 0.24
|  | 2024-04-15 |  | 2 | 7891 | 11258 | 13300 | 15574 | 18163 | 21031 | 15415 | 782.9961 | 13965 | 372 |  |  |  | -0.25 | -0.25 | -1.03 | 2.664 | 0.25
|  |  |  | 3 | 22075 | 31118 | 36366 | 42416 | 49590 | 53372 | 42413 | 810.7324 | 45227 | 474 |  |  |  | 0.5 | 0.5 | -0.01 | 1.048 | 0.5
| LV2071 | S3R9 |  | 1 | 896 | 1268 | 1572 | 1882 | 2220 | 2602 |  | 817.054 | 2085 | 251 |  |  |  | -0.72 | -0.72 | -1.42 | 12.038 | 0.72 |  | 797 | 1135 | 1416 | 1717 | 2035 | 2389 | 1693 | 830.1111 | 2110 | 261
|  |  |  | 2 | 6243 | 8972 | 10645 | 12546 | 14677 | 17127 |  | 817.332 | 13975 | 309 |  |  |  | -0.18 | -0.18 | -0.87 | 2.211 | 0.18 |  | 5533 | 7968 | 9448 | 11113 | 12982 | 15113 | 11017 | 837.088 | 14075 | 316
|  |  |  | 3 | 18381 | 26272 | 31017 | 36186 | 42060 | 48922 |  | 835.7415 | 44883 | 430 |  |  |  | -0.26 | -0.26 |  | 0.958 | 0.26 |  | 16010 | 22859 | 27077 | 31887 | 37124 | 42877 | 31793 | 857.751 |  |  | 850 | 42948 | 422
| LV2175 | S3R8 | 12:12PM->2:21PM | 1 | 1055 | 1634 | 1963 | 2337 | 2756 | 3226 | 2291 | 784.1762 | 2096 | 257 |  |  |  | -0.19 | -0.19 | -2.01 | 12.261 | 0.19
|  |  |  | 2 | 7657 | 10879 | 12800 | 14960 | 17223 | 20146 | 14941 | 789.2516 | 14038 | 309 |  |  |  | 0.27 | 0.27 | -0.13 | 2.201 | 0.27
|  |  |  | 3 | 21482 | 29961 | 35000 | 40464 | 46560 | 53472 | 40470 | 818.5187 | 44797 | 397 |  |  |  | -0.45 | -0.45 | 0.01 | 0.886 | 0.45
| LV2034 | S3R7 | 12:12PM->2:42PM | 1 | 1031 | 1532 | 1843 | 2188 | 2570 | 2998 | 2122 | 794.5062 | 2097 | 283 |  |  |  | -0.14 | -0.14 | -3.11 | 13.495 | 0.14
|  |  |  | 2 | 7458 | 10515 | 12357 | 14420 | 16746 | 19343 | 14260 | 795.0824 | 13971 | 323 |  |  |  | -0.21 | -0.21 | -1.12 | 2.312 | 0.21
|  |  |  | 3 | 21192 | 29397 | 34208 | 39412 | 45693 | 52700 | 39407 | 822.0363 | 44898 | 418 |  |  |  | -0.23 | -0.23 | -0.01 | 0.931 | 0.23
| LV2081 | S3R6 | 12:12PM->3:14PM | 1 | 1029 | 1581 | 1880 | 2214 | 2573 | 2988 | 2168 | 791.6095 | 2090 | 258 |  |  |  | -0.48 | -0.48 | -2.12 | 12.344 | 0.48
|  |  |  | 2 | 7400 | 10385 | 12174 | 14158 | 16384 | 18895 | 14003 | 798.3024 | 13889 | 301 |  |  |  | -0.8 | -0.8 | -1.11 | 2.167 | 0.8
|  |  |  | 3 | 21109 | 29246 | 33880 | 39474 | 45712 | 52655 | 39414 | 822.3212 | 44954 | 411 |  |  |  | -0.1 | -0.1 | -0.15 | 0.914 | 0.1
| LV2152 | S3R5 |  | 1 | 1086 | 1595 | 1876 | 2193 | 2544 | 2945 | 2175 | 792.6338 | 2085 | 275 |  |  |  | -0.72 | -0.72 | -0.83 | 13.189 | 0.72
|  |  |  | 2 | 8321 | 11519 | 13416 | 15534 | 17874 | 20475 | 15349 | 781.035 | 13781 | 324 |  |  |  | -1.59 | -1.59 | -1.21 | 2.351 | 1.59
|  |  |  | 3 | 23244 | 31263 | 35883 | 40928 | 46390 | 52298 | 40700 | 818.5378 | 44784 | 395 |  |  |  | -0.48 | -0.48 | -0.56 | 0.882 | 0.48
| LV2160 | S3R4 | 8:03AM->10:32AM | 1 | 1357 | 1888 | 2185 | 2522 | 2882 | 3284 | 2431 | 767.0389 | 2082 | 300 |  |  |  | -0.86 | -0.86 | -3.74 | 14.409 | 0.86
|  |  |  | 2 | 9427 | 12579 | 14368 | 16317 | 18417 | 20640 | 16055 | 769.0747 | 13873 | 346 |  |  |  | -0.92 | -0.92 | -1.63 | 2.494 | 0.92
|  |  |  | 3 | 25305 | 33121 | 37516 | 42194 | 47194 | 52413 | 42120 | 813.6045 | 44962 | 423 |  |  |  | -0.08 | -0.08 | -0.18 | 0.941 | 0.08
| LV2089 | S3R3 | 8:03AM->10:06AM | 1 | 1113 | 1723 | 2062 | 2440 | 2869 | 3341 | 2397 | 777.4189 | 2103 | 274 |  |  |  | 0.14 | 0.14 | -1.79 | 13.029 | 0.14
|  |  |  | 2 | 7834 | 11047 | 12973 | 15147 | 17558 | 20258 | 15074 | 787.177 | 13922 | 322 |  |  |  | -0.56 | -0.56 | -0.48 | 2.313 | 0.56
|  |  |  | 3 | 21705 | 30402 | 35669 | 41312 | 47925 | 55135 | 41283 | 813.6591 | 44637 | 424 | 814.5 | 44957 | 422 | -0.81 | -0.1 | -0.07 | 0.95 | 0.1
| LV2178 | S3R2 | 8:03AM->9:42AM | 1 | 1191 | 1807 | 2151 | 2544 | 2989 | 3483 | 2582 | 771.291 | 2101 | 284 |  |  |  | 0.05 | 0.05 | 1.47 | 13.517 | 0.05
|  |  |  | 2 | 8601 | 11960 | 13941 | 16181 | 18678 | 21446 | 16259 | 775.4154 | 13991 | 334 |  |  |  | -0.06 | -0.06 | 0.48 | 2.387 | 0.06
|  |  |  | 3 | 23569 | 31918 | 36647 | 42045 | 48114 | 54646 | 42002 | 812.7645 | 44978 | 405 |  |  |  | -0.05 | -0.05 | -0.1 | 0.9 | 0.05
| LV2096 | S3R1 | 8:03AM->9:00AM | 1 | 1802 | 2683 | 3222 | 3831 | 4508 | 5262 | 3909 | 718.2757 | 2105 | 316 |  |  |  | 0.24 | 0.24 | 2 | 15.012 | 0.24
|  |  |  | 2 | 12112 | 16861 | 19638 | 22702 | 26088 | 29556 | 22849 | 720.9477 | 13966 | 358 |  |  |  | -0.24 | -0.24 | 0.64 | 2.563 | 0.24
|  |  |  | 3 | 31215 | 42257 | 48633 | 55540 | 61913 | saturates | 55659 | 760.8613 | 44934 | 425 |  |  |  | -0.15 | -0.15 | 0.21 | 0.946 | 0.15
S2R | LV2130 | S2R10 | 9:05AM->10:10AM | 1 | 1662 | 2324 | 2716 | 3160 | 3651 | 4185 | 3194 | 734.2521 | 2093 | 287 |  |  |  | -0.33 | -0.33 | 1.06 | 13.712 | 0.33
|  |  |  | 2 | 10832 | 14616 | 16798 | 19159 | 21722 | 24468 | 18989 | 742.3457 | 13931 | 340 |  |  |  | -0.5 | -0.5 | -0.9 | 2.441 | 0.5 |  | *LV2093 has low sensitivity not enough
|  |  |  | 3 | 28626 | 37707 | 42882 | 48496 | 54438 | 60067 | 47877 | 784.6497 | 44919 | 433 |  |  |  | -0.18 | -0.18 | -1.29 | 0.964 | 0.18 |  | *LV2093 at 750 V has really distinct second peak
| LV2076 | S2R9 | 9:05AM->10:34AM | 1 | 1635 | 2337 | 2818 | 3326 | 3938 | 4576 | 3361 | 733.7856 | 2066 |  | 735 | 2116 | 278 | -1.65 | 0.76 | 1.04 | 0 | 0.76
|  |  |  | 2 | 10881 | 15034 | 17498 | 20154 | 23029 | 26214 | 19977 | 738.1152 | 13960 | 310 |  |  |  | -0.29 | -0.29 | -0.89 | 2.221 | 0.29
|  |  |  | 3 | 27990 | 37510 | 43019 | 48979 | 55154 | 60298 | 48704 | 783.2959 | 44927 | 370 |  |  |  | -0.16 | -0.16 | -0.56 | 0.824 | 0.16
| LV2172 | S2R8 | 9:05AM->11:19AM | 1 | 1440 | 2100 | 2471 | 2844 | 3356 | 3951 | 2920 | 751.8375 | 2119 | 225 | 750 | 2088 | 237 | 0.9 | -0.57 | 2.6 | 10.618 | 0.57
|  |  |  | 2 | 9672 | 13467 | 15743 | 18276 | 21175 | 24376 | 18223 | 756.0016 | 13965 | 300 |  |  |  | -0.25 | -0.25 | -0.29 | 2.148 | 0.25
|  |  |  | 3 | 26279 | 35715 | 41006 | 47157 | 53780 | 59497 | 46931 | 790.8376 | 44733 | 406 |  |  |  | -0.6 | -0.6 | -0.48 | 0.908 | 0.6
| LV2168 | S2R7 | 9:05AM->11:50AM | 1 | 1294 | 1851 | 2212 | 2462 | 2854 | 3291 | 2482 | 769.6258 | 2091 | 328 |  |  |  | -0.43 | -0.43 | 0.81 | 15.686 | 0.43
|  |  |  | 2 | 9323 | 12501 | 14348 | 16352 | 18516 | 20818 | 16171 | 770.1215 | 13892 | 329 |  |  |  | -0.78 | -0.78 | -1.12 | 2.368 | 0.78
|  |  |  | 3 | 25606 | 33808 | 38457 | 43512 | 48974 | 54192 | 43288 | 807.2204 | 44880 | 430 |  |  |  | -0.27 | -0.27 | -0.52 | 0.958 | 0.27
| LV2185 | S2R6 | 9:05AM->12:17PM | 1 | 1239 | 1749 | 1985 | 2281 | 2640 | 2922 | 2026 | 782.8632 | 2084 | 250 |  |  |  | -0.77 | -0.77 | -12.59 | 11.996 | 0.77
|  |  |  | 2 | 8537 | 11346 | 13027 | 14782 | 16698 | 18748 | 13033 | 789.2629 | 14028 | 302 |  |  |  | 0.2 | 0.2 | -13.42 | 2.153 | 0.2
|  |  |  | 3 | 23470 | 30823 | 34962 | 39494 | 44300 | 49426 | 35261 | 828.5296 | 44995 | 433 |  |  |  | -0.01 | -0.01 | -12 | 0.962 | 0.01
| LV2158 | S2R5 | 1:12PM->3:50PM | 1 | 1688 | 2276 | 2679 | 3064 | 3595 | 4082 | 3045 | 735.4921 | 2101 | 299 |  |  |  | 0.05 | 0.05 | -0.62 | 14.231 | 0.05
|  |  |  | 2 | 11279 | 15121 | 17273 | 19703 | 22328 | 25188 | 19721 | 736.4479 | 14000 | 301 |  |  |  | 0 | 0 | 0.09 | 2.15 | 0
|  |  |  | 3 | 30426 | 40328 | 45965 | 52040 | 57903 | 62401 | 52071 | 770.8811 | 44893 | 383 |  |  |  | -0.24 | -0.24 | 0.06 | 0.853 | 0.24
| LV2088 | S2R4 | 1:12PM->3:28PM | 1 | 961 | 1578 | 1850 | 2150 | 2570 | 2938 | 2200 | 792.6934 | 2013 | 303.5 |  |  |  | -4.32 | -4.32 | 2.27 | 15.077 | 4.32
|  |  |  | 2 | 7077 | 9937 | 11615 | 13569 | 15670 | 18104 | 13550 | 805.3002 | 13916 | 339 |  |  |  | -0.6 | -0.6 | -0.14 | 2.436 | 0.6
|  |  |  | 3 | 19484 | 27051 | 31594 | 36566 | 42114 | 48681 | 36598 | 836.2236 | 44879 | 404 |  |  |  | -0.27 | -0.27 | 0.09 | 0.9 | 0.27
| LV2181 | S2R3 | 1:12PM->3:04PM | 1 | 1550 | 2144 | 2513 | 2898 | 3291 | 3840 | 2937 | 746.5491 | 2067 | 274 | 748 | 2112 | 283 | -1.6 | 0.57 | 1.33 | 13.256 | 0.57
|  |  |  | 2 | 10537 | 14109 | 16130 | 18351 | 20766 | 23393 | 18417 | 748.6497 | 13967 | 322 |  |  |  | -0.24 | -0.24 | 0.36 | 2.305 | 0.24
|  |  |  | 3 | 28704 | 37936 | 43166 | 48821 | 54310 | 58538 | 48777 | 783.3765 | 45015 | 369 |  |  |  | 0.03 | 0.03 | -0.09 | 0.82 | 0.03
| LV2118 | S2R2 | 1:12PM->2:44PM | 1 | 1667 | 2275 | 2684 | 3115 | 3588 | 4111 | 3144 | 735.7959 | 2083 | 305 |  |  |  | -0.82 | -0.82 | 0.92 | 14.642 | 0.82
|  |  |  | 2 | 11058 | 14904 | 17110 | 19563 | 22193 | 25065 | 19592 | 738.9841 | 13994 | 287 |  |  |  | -0.04 | -0.04 | 0.15 | 2.051 | 0.04
|  |  |  | 3 | 30164 | 40262 | 46094 | 52334 | 57390 | 61507 | 52418 | 770.4919 | 44839 | 394 |  |  |  | -0.36 | -0.36 | 0.16 | 0.879 | 0.36
| LV2195 | S2R1 | 1:12PM->2:16PM | 1 | 1497 | 2203 | 2699 | 3213 | 3765 | 4325 | 3239 | 740.9158 | 2072 | 313 |  |  |  | -1.35 | -1.35 | 0.8 | 15.106 | 1.35
|  |  |  | 2 | 10303 | 14608 | 17194 | 20096 | 23421 | 27112 | 20175 | 743.4787 | 13938 | 341 |  |  |  | -0.44 | -0.44 | 0.39 | 2.447 | 0.44
|  |  |  | 3 | 28303 | 39234 | 46043 | 53778 | 61360 | saturates | 53714 | 770.3869 | 44483 | 397 | 772.5 | 45148 | 411 | -1.16 | 0.33 | -0.12 | 0.892 | 0.33
S4R | LV2124 | S4R10 | 4:18PM->5.00PM | 1 | 1719 | 2483 | 2951 | 3480 | 4084 | 4750 | 3461 | 726.0862 | 2084 | 301 |  |  |  | -0.77 | -0.77 | -0.55 | 14.443 | 0.77
|  |  | went back 18:45 | 2 | 11341 | 15635 | 18084 | 20780 | 23675 | 26660 | 20613 | 731.8332 | 13929 | 343 |  |  |  | -0.51 | -0.51 | -0.81 | 2.462 | 0.51
|  |  |  | 3 | 28580 | 38260 | 43797 | 49840 | 56325 | 62203 | 49750 | 779.691 | 44911 | 421 |  |  |  | -0.2 | -0.2 | -0.18 | 0.937 | 0.2
| LV2074 | S4R9 |  | 1 | 1058 | 1061 | 1915 | 2271 | 2670 | 3098 | 2291 | 787.6151 | 2100 | 264 |  |  |  | 0 | 0 | 0.87 | 12.571 | 0
|  |  | went back 18:50 | 2 | 7345 | 10408 | 12261 | 14318 | 16610 | 19200 | 14250 | 796.7174 | 14050 | 313 |  |  |  | 0.36 | 0.36 | -0.48 | 2.228 | 0.36
|  |  |  | 3 | 20732 | 29062 | 33855 | 38909 | 45004 | 51698 | 38810 | 824.3087 | 44840 | 404 | 825 | 44933 | 406 | -0.36 | -0.15 | -0.26 | 0.901 | 0.15
| LV2092 | S4R8 | 4:18PM->5.50PM | 1 | 1207 | 1835 | 2187 | 2593 | 3072 | 3562 | 2548 | 768.135 | 2097 | 262 |  |  |  | -0.14 | -0.14 | -1.77 | 12.494 | 0.14
|  |  | went back 18:55 | 2 | 8150 | 11605 | 13709 | 15972 | 18630 | 21646 | 15900 | 779.0444 | 13959 | 310 |  |  |  | -0.29 | -0.29 | -0.45 | 2.221 | 0.29
|  |  |  | 3 | 22744 | 32147 | 37635 | 43996 | 51500 | 58391 | 44016 | 803.4289 | 44953 | 407 |  |  |  | -0.1 | -0.1 | 0.05 | 0.905 | 0.1
| LV2073 | S4R7 | 4:18PM->6.10PM | 1 | 1488 | 2147 | 2525 | 2947 | 3426 | 3939 | 2934 | 746.5708 | 2096 | 287 |  |  |  | -0.19 | -0.19 | -0.44 | 13.693 | 0.19
|  |  |  | 2 | 9731 | 13188 | 15159 | 17327 | 19652 | 22137 | 17296 | 760.5404 | 14022 | 318 |  |  |  | 0.16 | 0.16 | -0.18 | 2.268 | 0.16
|  |  |  | 3 | 24562 | 32420 | 36888 | 41710 | 46960 | 52592 | 41700 | 815.9476 | 45025 | 389 |  |  |  | 0.06 | 0.06 | -0.02 | 0.864 | 0.06
| LV2191 | S4R6 | 4:18PM->6.30PM | 1 | 1452 | 2008 | 2337 | 2700 | 3065 | 3513 | 2710 | 757.3347 | 2087 | 257 |  |  |  | -0.62 | -0.62 | 0.37 | 12.314 | 0.62
|  |  |  | 2 | 10408 | 13860 | 15867 | 18038 | 20384 | 22934 | 18005 | 751.1624 | 13925 | 298 |  |  |  | -0.54 | -0.54 | -0.18 | 2.14 | 0.54
|  |  |  | 3 | 28403 | 37306 | 42350 | 47780 | 53497 | 58258 | 47750 | 787.239 | 44982 | 395 |  |  |  | -0.04 | -0.04 | -0.06 | 0.878 | 0.04
| LV2188 | S4R5 |  | 1 | out of range | 1496 | 1761 | 2044 | 2363 | 2721 | 2006 | 804.3156 | 2093 | 255 |  |  |  | -0.33 | -0.33 | -1.89 | 12.183 | 0.33
|  |  |  | 2 | 7899 | 10796 | 12521 | 14441 | 16551 | 18858 | 14292 | 794.7037 | 13963 | 297 |  |  |  | -0.26 | -0.26 | -1.04 | 2.127 | 0.26
|  |  | 10:42 | 3 | 22165 | 29450 | 33615 | 38147 | 43031 | 48298 | 38053 | 834.2369 | 44679 | 377 | 836 | 45073 | 375 | -0.72 | 0.16 | -0.25 | 0.844 | 0.16
| LV2184 | S4R4 |  | 1 | 1229 | 1858 | 2208 | 2598 | 3042 | 3526 | 2559 | 766.8609 | 2082 | 285 |  |  |  | -0.86 | -0.86 | -1.52 | 13.689 | 0.86
|  |  |  | 2 | 8884 | 12238 | 14190 | 16364 | 18756 | 21420 | 16259 | 772.1272 | 13869 | 319 |  |  |  | -0.94 | -0.94 | -0.65 | 2.3 | 0.94
|  |  |  | 3 | 23940 | 31849 | 36196 | 41017 | 46432 | 52312 | 41073 | 817.4398 | 44540 | 379 | 819 | 44912 | 379 | -1.03 | -0.2 | 0.14 | 0.851 | 0.2
| LV2137 | S4R3 |  | 1 | 1172 | 1718 | 2017 | 2357 | 2737 | 3163 | 2332 | 780.6613 | 2087 | 261 |  |  |  | -0.62 | -0.62 | -1.07 | 12.506 | 0.62
|  |  |  | 2 | 8784 | 12143 | 14121 | 16302 | 18692 | 21317 | 16291 | 773.148 | 13922 | 302 |  |  |  | -0.56 | -0.56 | -0.07 | 2.169 | 0.56
|  |  |  | 3 | 24236 | 32411 | 36990 | 41917 | 47164 | 52713 | 42061 | 814.35 | 44801 | 367 |  |  |  | -0.44 | -0.44 | 0.34 | 0.819 | 0.44
| LV2057 | S4R2 |  | 1 | 1735 | 2572 | 3085 | 3675 | 4328 | 5063 | 3601 | 722.3549 | 2078 | 271 |  |  |  | -1.06 | -1.06 | -2.05 | 13.041 | 1.06
|  |  |  | 2 | 11897 | 16764 | 19636 | 22847 | 26377 | 30146 | 22870 | 722.5242 | 13957 | 311 |  |  |  | -0.31 | -0.31 | 0.1 | 2.228 | 0.31
|  |  |  | 3 | 31777 | 43401 | 50125 | 56704 | 62066 | saturates | 56765 | 762.0552 | 46489 | 396 | 757 | 45179 | 385 | 3.2 | 0.4 | 0.11 | 0.852 | 0.4 |  | pretty bad fit
| LV2127 | S4R1 | 7:50:00 AM ->8:50 AM | 1 | 1662 | 2436 | 2910 | 3447 | 4052 | 4723 | 3494 | 729.6198 | 2081 | 299 |  |  |  | -0.91 | -0.91 | 1.35 | 14.368 | 0.91
|  |  |  | 2 | 11217 | 15645 | 18287 | 21299 | 24676 | 28465 | 21365 | 731.9464 | 13876 | 337 |  |  |  | -0.89 | -0.89 | 0.31 | 2.429 | 0.89
|  |  |  | 3 | 29976 | 40751 | 47311 | 54711 | 61939 | saturates | 54891 | 766.5932 | 44880 | 415 |  |  |  | -0.27 | -0.27 | 0.33 | 0.925 | 0.27
[18 empty] | σ: | 0.6291 | % | Average Gain Accuracy: | 0.4631
Note on S1: Dziubinski, Sean: % error = (experimental - actual) / actual *100
Note on M7: ====== ID#AAABLtEvVP8 Sean Dziubinski (2024-04-13 15:25:39) fit without point at 825 because it was weird 2 total reactions Sean Dziubinski reacted with 😁 at 2024-04-15 12:57 PM Sean Dziubinski reacted with 🤪 at 2024-04-13 08:35 AM

## LED2 Fitting v2
PMT # | LV2143 | LV2177 | LV1762 | LV2085 | LV2138 | LV2125 | LV2148 | LV2060 | LV2029 | LV2050 | LV2119 | LV2086 | LV2193 | LV2182 | LV2134 | LV2201 | LV2032 | LV2070 | LV1788 | LV2121 | LV2122 | LV2090 | LV2062 | LV2063 | LV2126 | LV2136 | LV2176 | LV2079 | LV2083 | LV2200 | LV2146 | LV2155 | LV2169 | LV2084 | LV2199 | LV2144 | LV2065 | LV2031 | LV2167 | LV2128 | LV2194 | LV2140 | LV2075 | LV2131 | LV2192 | LV2165 | LV2087 | LV2064 | LV2117 | LV2098 | LV2100 | LV1785 | LV2091 | LV2095 | LV2187 | LV2190 | LV2186 | LV2077 | LV1760 | LV2123 | LV1780 | LV1789 | LV2179 | LV2173 | LV2166 | LV2157 | LV2082 | LV2072 | LV2058 | LV2161 | LV2093 | LV2052 | LV2078 | LV2174 | LV2164 | LV2030 | LV2163 | LV2066 | LV2067 | LV2094 | LV2106 | LV2145 | LV2069 | LV2171 | LV2147 | LV2159 | LV2162 | LV2054 | LV2047 | LV2139 | LV2120 | LV2071 | LV2175 | LV2034 | LV2081 | LV2152 | LV2160 | LV2089 | LV2178 | LV2096 | LV2130 | LV2076 | LV2172 | LV2168 | LV2185 | LV2158 | LV2088 | LV2181 | LV2118 | LV2195 | LV2124 | LV2074 | LV2092 | LV2073 | LV2191 | LV2188 | LV2184 | LV2137 | LV2057 | LV2127 | 0 | 0 | 0 | 0
700 | 9375 | 13226 | 9769 | 8238 | 10904 | 9309 | 6938 | 8217 | 12849 | 8680 | 15053 | 7965 | 6748 | 8666 | 10810 | 8409 | 9414 | 7428 | 7764 | 9869 | 8004 | 7771 | 9784 | 11776 | 9064 | 8333 | 8674 | 5717 | 10056 | 13261 | 9856 | 8943 | 7805 | 7525 | 7487 | 8166 | 9641 | 9838 | 11560 | 8824 | 9655 | 12431 | 10639 | 12561 | 9687 | 10481 | 15245 | 9065 | 12490 | 11418 | 11140 | 7049 | 8425 | 9121 | 8621 | 8993 | 0 | 8838 | 11686 | 11434 | 10664 | 9943 | 8975 | 8556 | 9170 | 10077 | 9434 | 6906 | 9598 | 7521 | 9140 | 9430 | 8001 | 9784 | 9419 | 8165 | 11293 | 8937 | 9093 | 10006 | 11176 | 9823 | 9557 | 10387 | 9291 | 7951 | 10988 | 8984 | 9230 | 12178 | 7891 | 6243 | 7657 | 7458 | 7400 | 8321 | 9427 | 7834 | 8601 | 12112 | 10832 | 10881 | 9672 | 9323 | 8537 | 11279 | 7077 | 10537 | 11058 | 10303 | 11341 | 7345 | 8150 | 9731 | 10408 | 7899 | 8884 | 8784 | 11897 | 11217 | 0 | 0 | 0 | 0
750 | 13167 | 17951 | 13642 | 11726 | 15497 | 13262 | 9933 | 11749 | 17879 | 12427 | 20211 | 11328 | 9231 | 12193 | 14338 | 11963 | 12714 | 10168 | 10729 | 13249 | 11050 | 11151 | 14041 | 16821 | 12828 | 11436 | 11970 | 8071 | 14459 | 19139 | 13753 | 11973 | 10964 | 10626 | 10501 | 11611 | 13735 | 14102 | 15661 | 12589 | 13677 | 16772 | 14467 | 18114 | 13304 | 14992 | 17965 | 12958 | 16940 | 15681 | 15479 | 10078 | 12020 | 13165 | 11977 | 12327 | 10385 | 12247 | 16207 | 15653 | 14876 | 14391 | 12512 | 11715 | 12653 | 13843 | 12773 | 9807 | 13828 | 10680 | 13073 | 13445 | 11419 | 13080 | 13094 | 11594 | 15521 | 12726 | 12782 | 13613 | 16003 | 13813 | 13377 | 14223 | 12755 | 10899 | 15120 | 12850 | 13224 | 17317 | 11258 | 8972 | 10879 | 10515 | 10385 | 11519 | 12579 | 11047 | 11960 | 16861 | 14616 | 15034 | 13467 | 12501 | 11346 | 15121 | 9937 | 14109 | 14904 | 14608 | 15635 | 10408 | 11605 | 13188 | 13860 | 10796 | 12238 | 12143 | 16764 | 15645 | 0 | 0 | 0 | 0
775 | 15445 | 20644 | 15939 | 13857 | 18265 | 15590 | 11703 | 13860 | 20827 | 14652 | 23099 | 13365 | 10675 | 14332 | 16332 | 14067 | 14585 | 11769 | 12441 | 15208 | 12839 | 13202 | 16668 | 19893 | 15112 | 13260 | 13881 | 9503 | 17188 | 22770 | 16045 | 13705 | 12835 | 12477 | 12290 | 13695 | 16230 | 16739 | 18025 | 14875 | 16076 | 19263 | 16678 | 21729 | 15430 | 17784 | 20892 | 15326 | 19500 | 18182 | 17988 | 11916 | 14222 | 15659 | 13979 | 14271 | 12439 | 14191 | 18861 | 18076 | 17380 | 17146 | 14615 | 13550 | 14698 | 16038 | 14000 | 11572 | 16380 | 12675 | 15454 | 15899 | 13474 | 14940 | 15282 | 13607 | 18002 | 15018 | 14946 | 16658 | 18933 | 16184 | 15672 | 16458 | 14771 | 12618 | 17545 | 15182 | 15677 | 20503 | 13300 | 10645 | 12800 | 12357 | 12174 | 13416 | 14368 | 12973 | 13941 | 19638 | 16798 | 17498 | 15743 | 14348 | 13027 | 17273 | 11615 | 16130 | 17110 | 17194 | 18084 | 12261 | 13709 | 15159 | 15867 | 12521 | 14190 | 14121 | 19636 | 18287 | 0 | 0 | 0 | 0
800 | 18005 | 23539 | 18537 | 16276 | 21306 | 18223 | 13712 | 16272 | 24079 | 17305 | 26152 | 15664 | 12279 | 16723 | 18455 | 16463 | 16626 | 13511 | 14322 | 17357 | 14830 | 15532 | 19649 | 23329 | 17642 | 15297 | 16009 | 11129 | 20330 | 27965 | 19624 | 15594 | 14955 | 14564 | 14320 | 16043 | 19080 | 19716 | 20596 | 17463 | 18788 | 22022 | 19039 | 25629 | 17798 | 21031 | 24074 | 18029 | 22290 | 20945 | 20784 | 13969 | 16724 | 18534 | 16225 | 16432 | 14734 | 16368 | 21792 | 20728 | 20157 | 20298 | 16944 | 15600 | 16990 | 18483 | 16753 | 13527 | 19353 | 14885 | 18183 | 18706 | 15769 | 16999 | 17732 | 15973 | 20777 | 17640 | 17353 | 17829 | 22286 | 18831 | 18138 | 18930 | 17006 | 14533 | 20217 | 17813 | 18425 | 24120 | 15574 | 12546 | 14960 | 14420 | 14158 | 15534 | 16317 | 15147 | 16181 | 22702 | 19159 | 20154 | 18276 | 16352 | 14782 | 19703 | 13569 | 18351 | 19563 | 20096 | 20780 | 14318 | 15972 | 17327 | 18038 | 14441 | 16364 | 16302 | 22847 | 21299 | 0 | 0 | 0 | 0
825 | 20865 | 26550 | 21416 | 19032 | 24811 | 21282 | 15962 | 18988 | 27610 | 20259 | 29627 | 18245 | 14045 | 19381 | 20744 | 19129 | 18869 | 15416 | 16401 | 19674 | 17015 | 18180 | 22979 | 27128 | 20507 | 17526 | 18337 | 12959 | 23942 | 31707 | 21423 | 17615 | 17311 | 16914 | 16616 | 18686 | 22308 | 23115 | 23343 | 20389 | 21844 | 25064 | 21606 | 30021 | 20382 | 24787 | 27351 | 21117 | 25306 | 23990 | 23802 | 16278 | 19581 | 21859 | 18729 | 18788 | 17273 | 18733 | 24950 | 23534 | 23270 | 23875 | 19542 | 17855 | 19495 | 21211 | 18979 | 15737 | 22705 | 17268 | 21305 | 21859 | 18334 | 19203 | 20454 | 18585 | 23813 | 20626 | 19968 | 20109 | 26067 | 21806 | 20849 | 21640 | 19469 | 16599 | 23179 | 20772 | 21653 | 28149 | 18163 | 14677 | 17223 | 16746 | 16384 | 17874 | 18417 | 17558 | 18678 | 26088 | 21722 | 23029 | 21175 | 18516 | 16698 | 22328 | 15670 | 20766 | 22193 | 23421 | 23675 | 16610 | 18630 | 19652 | 20384 | 16551 | 18756 | 18692 | 26377 | 24676 | 0 | 0 | 0 | 0
850 | 24082 | 29654 | 24641 | 22131 | 28823 | 24590 | 18512 | 22029 | 31321 | 23697 | 33384 | 21193 | 15990 | 22383 | 23181 | 22089 | 21224 | 17468 | 18621 | 22198 | 19431 | 21178 | 26498 | 31153 | 23696 | 19982 | 20928 | 15001 | 28120 | 36887 | 24511 | 19795 | 19938 | 19580 | 19159 | 21677 | 25914 | 26961 | 26353 | 23696 | 25246 | 28358 | 24371 | 34608 | 23245 | 28893 | 0 | 24637 | 28573 | 27278 | 27070 | 18888 | 22799 | 25657 | 21512 | 21402 | 20036 | 21293 | 28366 | 26498 | 26709 | 27923 | 22440 | 20287 | 22216 | 24217 | 21333 | 18214 | 26570 | 19957 | 24793 | 25450 | 21182 | 21587 | 23502 | 21518 | 27186 | 23807 | 22870 | 22575 | 30342 | 25107 | 23800 | 24628 | 22152 | 18896 | 26414 | 24079 | 25302 | 32435 | 21031 | 17127 | 20146 | 19343 | 18895 | 20475 | 20640 | 20258 | 21446 | 29556 | 24468 | 26214 | 24376 | 20818 | 18748 | 25188 | 18104 | 23393 | 25065 | 27112 | 26660 | 19200 | 21646 | 22137 | 22934 | 18858 | 21420 | 21317 | 30146 | 28465 | 0 | 0 | 0 | 0
t1 | 213 | 395 | 219 | 194 | 203 | 209.2 | 211 | 204 | 278.5 | 184 | 296 | 201 | 248 | 213 | 346 | 217 | 302 | 254 | 270 | 280 | 243 | 195 | 215 | 233.3 | 213 | 243 | 249 | 202 | 175 | 232 | 356 | 307 | 219 | 207 | 211 | 204 | 196 | 190 | 297 | 199 | 211 | 264 | 306 | 220 | 245 | 181 | 389 | 189 | 291 | 254 | 270 | 208 | 192 | 177 | 221 | 248 | 254.4 | 268 | 266 | 316 | 228 | 186 | 228 | 253 | 245 | 238 | 203 | 211 | 185 | 213 | 192 | 192 | 218.5 | 298 | 221 | 204 | 239 | 207 | 247 | 3032 | 194 | 221 | 258 | 252 | 252 | 255.5 | 248 | 207 | 186 | 210 | 209.5 | 194 | 207 | 211 | 218 | 229 | 315 | 215 | 222 | 275 | 297 | 270 | 212 | 298 | 302 | 276 | 214 | 289 | 279 | 205.5 | 316 | 216 | 207 | 300 | 287 | 237.5 | 242 | 249 | 247 | 206
A1 | 14382 | 36573 | 15237 | 11998 | 16327 | 15033 | 11175 | 12871 | 26582 | 11977 | 27944 | 11933 | 11232 | 13378 | 22901 | 13910 | 18945 | 12413 | 14954 | 17445 | 13432 | 11647 | 17107 | 22002 | 14334 | 13751 | 14817 | 8515 | 13381 | 16160 | 6860 | 17341 | 12433 | 11374 | 11421 | 12445 | 14321 | 14318 | 22347 | 13266 | 15227 | 21122 | 21752 | 22997 | 16069 | 14727 | 41181 | 12898 | 24010 | 20055 | 21719 | 11153 | 12187 | 12468 | 13338 | 15058 | 19986 | 16981 | 22403 | 25562 | 17313 | 14625 | 14458 | 14791 | 15775 | 16221 | 4109 | 10884 | 13666 | 12108 | 13486 | 13625 | 13443 | 18297 | 14583 | 12524 | 18294 | 14483 | 16548 | 88341 | 16655 | 15890 | 18146 | 17606 | 16057 | 13640 | 18723 | 14343 | 13023 | 20124 | 12640 | 9272 | 10972 | 11536 | 11541 | 13256 | 18873 | 12380 | 13428 | 25373 | 20995 | 20298 | 14443 | 17825 | 15750 | 19574 | 10769 | 18997 | 19746 | 15810 | 26538 | 11738 | 12453 | 19546 | 18325 | 12606 | 14662 | 15329 | 22475 | 16336
x0 | 700 | 709 | 700 | 701 | 699 | 706.3 | 700.5 | 701.8 | 706.2 | 701 | 702 | 698 | 700 | 698 | 701 | 702 | 708 | 700 | 705 | 700 | 700.2 | 700.6 | 705 | 704.4 | 700 | 700 | 699 | 701 | 700 | 588 | 195 | 700 | 701 | 700 | 703 | 699 | 701.5 | 700 | 697 | 700 | 701 | 703 | 699 | 701 | 699 | 704 | 707 | 700 | 700 | 702 | 702.5 | 699 | 701 | 701.2 | 700.7 | 700.7 | 598.9 | 705 | 704 | 709 | 701 | 701 | 700 | 705 | 704 | 702 | 502 | 700 | 701 | 699 | 702 | 700.7 | 701 | 704 | 700.5 | 704 | 701.2 | 706.5 | 700.8 | -2425 | 701 | 701 | 699 | 700 | 702 | 698.6 | 701 | 701 | 700 | 706 | 701 | 698.6 | 687 | 700.6 | 699 | 700 | 706 | 701 | 702 | 713 | 703 | 695.3 | 702 | 702.6 | 696.6 | 704.2 | 698 | 701.7 | 700.9 | 702 | 715 | 699 | 697 | 706 | 700 | 702.6 | 700 | 701 | 706 | 701
y0 | -5028 | -22532 | -5441 | -3708 | -5509 | -5279 | -4210 | -4547 | -13173 | -3223 | -12675 | -4029 | -4441 | -4814 | -12032 | -5379 | -9026 | -4903 | -6929 | -7595 | -5423 | -3842 | -6967 | -9857 | -5269 | -5388 | -6164 | -2761 | -3323 | -13106 | -18575 | -8370 | -4564 | -3849 | -3789 | -4300 | -4579 | -4492 | -11008 | -4431 | -5453 | -8410 | -11183 | -10376 | -6414 | -3939 | -26195 | -3790 | -11495 | -8432 | -10390 | -4157 | -3710 | -3266 | -4677 | -6026 | -9642 | -7829 | -10415 | -13449 | -6583 | -4608 | -5495 | -5954 | -6342 | -6034 | -1432 | -4005 | -3979 | -4671 | -4192 | -4150 | -5384 | -8255 | -5136 | -4117 | -6914 | -5116 | -7407 | -237623 | -5373 | -5992 | -8631 | -7158 | -6598 | -5768 | -7613 | -5278 | -3793 | -7404 | -4695 | -3099 | -3981 | -4052 | -4200 | -4884 | -9057 | -4484 | -4691 | -12117 | -9969 | -9786 | -4633 | -8358 | -7406 | -8000 | -3780 | -8348 | -8634 | -5346 | -13990 | -4453 | -4510 | -9443 | -7901 | -4578 | -5753 | -6481 | -10030 | -4984
Fit | 759.626 | 708.5569 | 753.361 | 776.5196 | 735.1453 | 758.3424 | 803.5296 | 776.3276 | 712.3241 | 767.8394 | 688.2432 | 780.9475 | 822.9608 | 770.6308 | 745.3385 | 773.9533 | 766.9154 | 806.8264 | 795.7624 | 759.7544 | 789.8228 | 783.7688 | 748.7443 | 723.2843 | 763.0189 | 783.4809 | 775.7228 | 837.7995 | 745.1849 | 707.9941 | 749.5902 | 778.1767 | 788.7905 | 793.2778 | 796.4997 | 777.6588 | 752.5198 | 748.6059 | 730.4136 | 765.4371 | 752.6805 | 718.6267 | 743.8178 | 713.8118 | 757.6356 | 739.7103 | 697.5728 | 760.7756 | 717.4635 | 730.4505 | 733.8162 | 800.3683 | 772.7619 | 758.8266 | 775.1053 | 771.4103 | 641.6374 | 772.3067 | 726.8767 | 731.5064 | 740.4458 | 745.7998 | 768.1515 | 780.7509 | 766.294 | 752.2475 | 770.6225 | 806.208 | 751.7443 | 791.2514 | 759.4713 | 755.7587 | 780.9687 | 762.3572 | 760.549 | 779.3176 | 733.1891 | 763.9529 | 764.3908 | 748.6781 | 730.327 | 751.7509 | 755.9846 | 746.3121 | 764.7604 | 793.4053 | 736.5985 | 762.2104 | 758.0484 | 718.9496 | 782.9961 | 817.332 | 789.2516 | 795.0824 | 798.3024 | 781.035 | 769.0747 | 787.177 | 775.4154 | 720.9477 | 742.3457 | 738.1152 | 756.0016 | 770.1215 | 789.2629 | 736.4479 | 805.3002 | 748.6497 | 738.9841 | 743.4787 | 731.8332 | 796.7174 | 779.0444 | 760.5404 | 751.1624 | 794.7037 | 772.1272 | 773.148 | 722.5242 | 731.9464 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| S1T |  |  |  |  | S1B |  |  |  |  | S2T |  |  |  |  | S2B |  |  |  |  | S3T |  |  |  |  | S3B |  |  |  |  | S4T |  |  |  |  | S4B |  |  |  |  | S1L | [9 empty] | S1R | [9 empty] | S2L | [9 empty] | S4L | [9 empty] | S3L | [9 empty] | S3R | [9 empty] | S4L | [9 empty] | S4R
-7901.19627 ± 834.00818
700.35914 ± 9658902.99808
18325.63524 ± 6.15808E8
287.43606 ± 9.93231

## LED3 Fitting v2
PMT # | LV2143 | LV2177 | LV1762 | LV2085 | LV2138 | LV2125 | LV2148 | LV2060 | LV2029 | LV2050 | LV2119 | LV2086 | LV2193 | LV2182 | LV2134 | LV2201 | LV2032 | LV2070 | LV1788 | LV2121 | LV2122 | LV2090 | LV2062 | LV2063 | LV2126 | LV2136 | LV2176 | LV2079 | LV2083 | LV2200 | LV2146 | LV2155 | LV2169 | LV2084 | LV2199 | LV2144 | LV2065 | LV2031 | LV2167 | LV2128 | LV2194 | LV2140 | LV2075 | LV2131 | LV2192 | LV2165 | LV2087 | LV2064 | LV2117 | LV2098 | LV2100 | LV1785 | LV2091 | LV2095 | LV2187 | LV2190 | LV2186 | LV2077 | LV1760 | LV2123 | LV1780 | LV1789 | LV2179 | LV2173 | LV2166 | LV2157 | LV2082 | LV2072 | LV2058 | LV2161 | LV2093 | LV2052 | LV2078 | LV2174 | LV2164 | LV2030 | LV2163 | LV2066 | LV2067 | LV2094 | LV2106 | LV2145 | LV2069 | LV2171 | LV2147 | LV2159 | LV2162 | LV2054 | LV2047 | LV2139 | LV2120 | LV2071 | LV2175 | LV2034 | LV2081 | LV2152 | LV2160 | LV2089 | LV2178 | LV2096 | LV2130 | LV2076 | LV2172 | LV2168 | LV2185 | LV2158 | LV2088 | LV2181 | LV2118 | LV2195 | LV2124 | LV2074 | LV2092 | LV2073 | LV2191 | LV2188 | LV2184 | LV2137 | LV2057 | LV2127
700 | 27229 | 33658 | 29142 | 25983 | 32263 | 25883 | 19486 | 23808 | 33569 | 25413 | 41252 | 25215 | 21332 | 26554 | 33526 | 25276 | 27836 | 22700 | 22708 | 29673 | 23801 | 24443 | 31110 | 33089 | 27493 | 24940 | 25778 | 19544 | 32454 | 41050 | 29110 | 26520 | 24206 | 23656 | 23074 | 24863 | 30193 | 31270 | 33660 | 27288 | 25999 | 33575 | 27020 | 35721 | 25579 | 29776 | 38613 | 26026 | 32188 | 29191 | 28947 | 20049 | 23836 | 26842 | 23888 | 23852 | 0 | 23551 | 30518 | 29079 | 29003 | 28774 | 24850 | 23443 | 24503 | 26292 | 24085 | 19566 | 28041 | 19937 | 26354 | 27266 | 22807 | 26903 | 25738 | 22836 | 28816 | 25408 | 24767 | 29945 | 30702 | 26696 | 25935 | 27450 | 25069 | 21834 | 28782 | 25797 | 26353 | 33312 | 22075 | 18381 | 21482 | 21192 | 21109 | 23244 | 25305 | 21705 | 23569 | 31215 | 28626 | 27990 | 26279 | 25606 | 23470 | 30426 | 19484 | 28704 | 30164 | 28303 | 28580 | 20732 | 22744 | 24562 | 28403 | 22165 | 23940 | 24236 | 31777 | 29976
750 | 37560 | 44009 | 39313 | 36884 | 45005 | 35391 | 27302 | 33559 | 45086 | 35937 | 53964 | 35583 | 28345 | 36418 | 42849 | 35007 | 36546 | 30183 | 30148 | 39484 | 31611 | 34809 | 44114 | 45886 | 38277 | 33168 | 34213 | 27312 | 46597 | 57979 | 39151 | 34788 | 32987 | 33181 | 31390 | 34764 | 42667 | 44169 | 44472 | 37952 | 35665 | 44857 | 35355 | 51232 | 33983 | 42650 | 44708 | 37166 | 42638 | 38731 | 38883 | 28130 | 33885 | 38297 | 32148 | 31460 | 27706 | 31475 | 41011 | 38704 | 39002 | 41021 | 33760 | 31154 | 32488 | 35100 | 31660 | 27551 | 39772 | 27403 | 37184 | 38431 | 31857 | 35210 | 34568 | 31699 | 38126 | 36060 | 33468 | 34283 | 42984 | 36725 | 35136 | 36064 | 33230 | 28816 | 38270 | 35918 | 37315 | 46978 | 31118 | 26272 | 29961 | 29397 | 29246 | 31263 | 33121 | 30402 | 31918 | 42257 | 37707 | 37510 | 35715 | 33808 | 30823 | 40328 | 27051 | 37936 | 40262 | 39234 | 38260 | 29062 | 32147 | 32420 | 37306 | 29450 | 31849 | 32411 | 43401 | 40751
775 | 43775 | 49828 | 45195 | 43674 | 52762 | 41429 | 31975 | 39210 | 51560 | 42498 | 60521 | 41802 | 32295 | 42551 | 47872 | 40994 | 41440 | 34392 | 34213 | 45137 | 36217 | 41267 | 51946 | 52244 | 44840 | 37938 | 39070 | 31899 | 55313 | 65187 | 45011 | 39450 | 38039 | 38716 | 36230 | 40760 | 50268 | 51867 | 50577 | 44417 | 41771 | 51309 | 40287 | 59265 | 38901 | 50956 | 51421 | 43727 | 48614 | 44127 | 44662 | 32962 | 39879 | 45360 | 37039 | 35963 | 32634 | 35773 | 46958 | 44194 | 44817 | 48834 | 38714 | 35393 | 37087 | 40127 | 35900 | 32368 | 47110 | 31976 | 44176 | 45465 | 37023 | 39848 | 39965 | 37159 | 43570 | 42338 | 38544 | 39025 | 50718 | 42586 | 40312 | 41152 | 38048 | 32776 | 43725 | 41807 | 44373 | 55367 | 36366 | 31017 | 35000 | 34208 | 33880 | 35883 | 37516 | 35669 | 36647 | 48633 | 42882 | 43019 | 41006 | 38457 | 34962 | 45965 | 31594 | 43166 | 46094 | 46043 | 43797 | 33855 | 37635 | 36888 | 42350 | 33615 | 36196 | 36990 | 50125 | 47311
800 | 50738 | 55932 | 51658 | 51504 | 60851 | 48330 | 37137 | 45424 | 57521 | 50163 | Saturates | 49079 | 36526 | 49506 | 53195 | 47754 | 46695 | 38877 | 38421 | 51305 | 41259 | 48655 | 59742 | 59870 | 52257 | 43149 | 44380 | 36981 | 62851 | Saturates | 51349 | 44474 | 43553 | 45403 | 41863 | 47614 | 58237 | 59222 | 56993 | 51905 | 48732 | 57225 | 45641 | saturates | 44289 | 59845 | 58497 | 51708 | 55048 | 49982 | 50967 | 38281 | 46767 | 53423 | 42714 | 40884 | 38004 | 40504 | 53446 | 50150 | 51216 | 57086 | 44363 | 40010 | 42052 | 45605 | 40490 | 37610 | 54950 | 37200 | 52275 | 53322 | 42445 | 44856 | 45924 | 43296 | 49563 | 49738 | 44253 | 44188 | 59128 | 49529 | 46249 | 46675 | 43330 | 37152 | 49684 | 48549 | 52839 | 63363 | 42416 | 36186 | 40464 | 39412 | 39474 | 40928 | 42194 | 41312 | 42045 | 55540 | 48496 | 48979 | 47157 | 43512 | 39494 | 52040 | 36566 | 48821 | 52334 | 53778 | 49840 | 38909 | 43996 | 41710 | 47780 | 38147 | 41017 | 41917 | 56704 | 54711
825 | 57759 | 61256 | 57743 | 59218 | Saturates | 55919 | 42662 | 52513 | 62620 | 58065 | Saturates | 56919 | 41101 | 56788 | 58532 | 55126 | 52252 | 43627 | 43062 | 56695 | 46792 | 56342 | Saturates | 0 | 59410 | 48892 | 50157 | 42552 | Saturates | Saturates | 57971 | 49787 | 49707 | 52966 | 48207 | 55251 | 64817 | saturates | 62331 | 58921 | 56429 | 61705 | 51391 | saturates | 50103 | saturate | 64417 | 58746 | 60537 | 56182 | 57572 | 44049 | 54484 | 60386 | 49006 | 46241 | 43858 | 45723 | 59284 | 56605 | 57844 | 63693 | 50804 | 45247 | 47442 | 51567 | 45423 | 43376 | 61421 | 43070 | 60529 | 60121 | 48767 | 50149 | 52663 | 50200 | 55859 | 57007 | 50448 | 49763 | sat | 56667 | 52624 | 52673 | 49072 | 41903 | 55838 | 55413 | 59975 | saturates | 49590 | 42060 | 46560 | 45693 | 45712 | 46390 | 47194 | 47925 | 48114 | 61913 | 54438 | 55154 | 53780 | 48974 | 44300 | 57903 | 42114 | 54310 | 57390 | 61360 | 56325 | 45004 | 51500 | 46960 | 53497 | 43031 | 46432 | 47164 | 62066 | 61939
850 | 63571 | 0 | 62903 | Saturates | Saturates | 62685 | 49234 | 58678 | saturates | 64555 | Saturates | 63442 | 45967 | 62925 | 62886 | 61556 | 57992 | 48722 | 48069 | 60914 | 52652 | 62873 | Saturates | 0 | Saturates | 55197 | 56020 | 48582 | Saturates | Saturates | 63659 | 54788 | 55769 | 60208 | 55033 | 62113 | saturates | saturates | saturates | 64992 | 63205 | saturates | 56892 | saturates | 56207 | saturate | 0 | 64796 | saturates | 61719 | 63074 | 50863 | 60885 | saturates | 55805 | 52080 | 50151 | 51278 | 64292 | 62572 | 63561 | saturated | 57032 | 50956 | 53193 | 57665 | 50760 | 49993 | saturated | 50060 | saturated | saturates | 55249 | 55250 | 59178 | 57290 | 61394 | 63085 | 56834 | 55730 | sat | 62654 | 58354 | 58336 | 55245 | 47026 | 61108 | 61103 | saturates | saturates | 53372 | 48922 | 53472 | 52700 | 52655 | 52298 | 52413 | 55135 | 54646 | saturates | 60067 | 60298 | 59497 | 54192 | 49426 | 62401 | 48681 | 58538 | 61507 | saturates | 62203 | 51698 | 58391 | 52592 | 58258 | 48298 | 52312 | 52713 | saturates | saturates
t1 | 479.5 | 379 | 953 | 225 | 262 | 258.5 | 230 | 337 | 320 | 166.2 | 1204 | 317 | 338 | 345 | 575 | 302 | 389 | 368 | 388 | 519 | 269 | 335 | 201 | 137.5 | 266 | 263 | 315 | 252 | 179 | 233 | 512 | 461 | 308 | 226 | 215 | 272 | 589 | 212 | 1137 | 494 | 263 | 279 | 346 | 1066 | 291 | 193 | 335 | 575 | 665 | 492 | 524 | 227 | 305.5 | 311 | 217 | 260 | 305.6 | 324 | 299 | 395 | 488 | 154 | 285 | 288 | 302 | 320 | 327 | 229 | 167 | 182 | 181 | 275 | 303 | 440 | 254 | 223 | 408 | 393 | 267 | 121 | 162 | 347 | 353 | 335 | 265 | 287.5 | 509 | 427 | 212 | 182 | 525 | 209 | 236 | 206 | 201 | 286 | 375 | 223 | 241 | 406 | 408.6 | 548 | 370 | 385 | 332 | 305.5 | 213 | 308 | 295 | 237 | 407 | 243 | 241 | 309 | 590 | 301 | 279 | 358 | 257 | 245
A1 | 76119 | 73719 | 116890 | 37985 | 61451 | 38507 | 30984 | 50539 | 65339 | 29982 | 299918 | 50369 | 44630 | 52823 | 110282 | 45849 | 66003 | 51446 | 52859 | 72573 | 39536 | 54277 | 44555 | 25036 | 44613 | 39454 | 52358 | 36102 | 40723 | 63861 | 79469 | 63289 | 45408 | 34692 | 32797 | 42420 | 99721 | 46353 | 117456 | 66225 | 39658 | 54902 | 48062 | 323218 | 46569 | 44481 | 79251 | 75736 | 95138 | 74603 | 78878 | 31679 | 46389 | 50886 | 32982 | 36447 | 58130 | 47878 | 57712 | 61681 | 76207 | 31946 | 51943 | 40094 | 45266 | 55192 | 45944 | 32296 | 33743 | 23286 | 34611 | 57266 | 51753 | 61441 | 45584 | 37877 | 58812 | 61681 | 45643 | 9918 | 34151 | 51904 | 51359 | 47130 | 40439 | 38271 | 73094 | 63906 | 41981 | 43475 | 37082 | 28540 | 35375 | 29802 | 29426 | 42583 | 56438 | 34846 | 37019 | 85238 | 61003 | 77337 | 54506 | 52729 | 46069 | 55809 | 27658 | 52418 | 54964 | 40365 | 62843 | 35601 | 35902 | 45194 | 77955 | 40815 | 40326 | 55723 | 48417 | 48102
x0 | 564 | 700 | 179 | 662 | 700 | 645 | 691 | 626.4 | 700 | 700 | 700 | 621 | 702 | 611 | 737 | 632 | 710 | 697 | 695 | 532 | 704.4 | 616 | 700 | 700 | 650 | 699 | 715 | 702 | 700 | 700 | 567 | 626 | 666 | 671 | 707.5 | 647 | 464 | 700 | -159 | 455 | 642 | 700 | 648 | 700 | 705 | 700 | 700 | 374 | 449 | 594 | 551 | 692.3 | 624 | 605.7 | 705.5 | 701.3 | 600.5 | 705.6 | 700 | 633 | 579 | 700 | 730 | 699 | 703 | 715 | 700 | 695 | 700 | 698 | 700 | 700 | 706 | 640.7 | 720.7 | 710 | 604 | 584.6 | 714 | 689 | 700 | 607 | 632 | 646 | 703 | 703.6 | 561 | 576 | 700 | 700 | 188 | 696 | 696 | 703 | 706.4 | 702 | 707 | 699 | 707 | 700 | 635.3 | 535.8 | 621 | 646.7 | 703.5 | 700 | 694 | 700 | 700 | 658 | 621 | 696 | 664 | 702 | 526 | 702 | 702 | 705 | 700 | 700
y0 | -74046 | -40061 | -173025 | -19107 | -29189 | -21960.5 | -12733 | -39206 | -31770 | -4569 | -258666 | -39585 | -22971 | -41990 | -69812 | -32329 | -36456 | -29120 | -30812 | -70614 | -15131 | -45466 | -13445 | 8053 | -26386 | -14522 | -24159 | -16288 | -8270 | -22631 | -74070 | -47735 | -26500 | -15765 | -8784 | -26807 | -118796 | -15083 | -216305 | -81685 | -23651 | -21327 | -28829 | -287497 | -20136 | -14705 | -40638 | -107586 | -106611 | -63326 | -75994 | -12719 | -35873 | -42178 | -8299 | -12433 | -30327 | -23485 | -27194 | -44125 | -68817 | -3172 | -21963 | -16716 | -20275 | -26390 | -21793 | -13379 | -5703 | -3586 | -8257 | -30000 | -27952 | -43441 | -16389 | -13464 | -45722 | -57499 | -18590 | 18724 | -3449 | -41255 | -36250 | -27987 | -14816 | -14981 | -67229 | -59724 | -15628 | -10163 | -76424 | -10697 | -14451 | -8113 | -7384 | -19003 | -29995 | -13271 | -12414 | -54023 | -42921 | -76487 | -41257 | -35006 | -22137 | -25383 | -8927 | -23714 | -24800 | -19856 | -47809 | -15364 | -19029 | -20348 | -76378 | -18331 | -15993 | -30629 | -16640 | -18126
Fit | 778.4382 | 754.238 | 773.0777 | 779.7565 | 749.3545 | 788.0184 | 834.1423 | 798.4455 | 751.5919 | 783.5601 | 714.9528 | 785.3268 | 844.1881 | 783.102 | 760.1468 | 789.8601 | 791.8311 | 831.3762 | 834.9239 | 773.6797 | 817.1954 | 787.1426 | 754.5437 | 753.5108 | 775.0403 | 807.1484 | 802.6656 | 835.3674 | 748.0759 | 713.3643 | 774.0244 | 802.1179 | 805.8348 | 797.6741 | 813.8474 | 790.1705 | 756.2886 | 755.0015 | 750.1731 | 775.431 | 786.3194 | 752.7442 | 796.5239 | 730.1719 | 802.6428 | 756.8103 | 725.9655 | 776.773 | 758.8831 | 777.4985 | 775.1874 | 828.4847 | 793.8023 | 773.1313 | 809.6504 | 819.5376 | 679.6983 | 821.5786 | 766.9437 | 778.3853 | 774.7556 | 763.2526 | 802.3881 | 823.2192 | 813.548 | 797.3487 | 822.3552 | 830.5709 | 768.0044 | 831.8579 | 778.0033 | 774.1898 | 810.0257 | 800.9737 | 796.3092 | 806.797 | 780.8492 | 784.1957 | 802.5388 | 806.8909 | 756.6546 | 783.2455 | 793.9178 | 792.5195 | 806.7418 | 832.7839 | 779.2567 | 786.9017 | 777.9184 | 743.3354 | 810.7324 | 835.7415 | 818.5187 | 822.0363 | 822.3212 | 818.5378 | 813.6045 | 813.6591 | 812.7645 | 760.8613 | 784.6497 | 783.2959 | 790.8376 | 807.2204 | 828.5296 | 770.8811 | 836.2236 | 783.3765 | 770.4919 | 770.3869 | 779.691 | 824.3087 | 803.4289 | 815.9476 | 787.239 | 834.2369 | 817.4398 | 814.35 | 762.0552 | 766.5932
| S1T |  |  |  |  | S1B |  |  |  |  | S2T |  |  |  |  | S2B |  |  |  |  | S3T |  |  |  |  | S3B |  |  |  |  | S4T |  |  |  |  | S4B |  |  |  |  | S1L | [9 empty] | S1R | [9 empty] | S2L | [9 empty] | S4L | [9 empty] | S3L | [9 empty] | S3R | [9 empty] | S4L | [9 empty] | S4R
-76378.39225 ± 89173.92065
526.12986 ± 8.5477E7
77955.38515 ± 1.12896E10
590.22504 ± 440.71863
scrath work:

## LED1 Fitting v2
PMT # | LV2143 | LV2177 | LV1762 | LV2085 | LV2138 | LV2125 | LV2148 | LV2060 | LV2029 | LV2050 | LV2119 | LV2086 | LV2193 | LV2182 | LV2134 | LV2201 | LV2032 | LV2070 | LV1788 | LV2121 | LV2122 | LV2090 | LV2062 | LV2063 | LV2126 | LV2136 | LV2176 | LV2079 | LV2083 | LV2200 | LV2146 | LV2155 | LV2169 | LV2084 | LV2199 | LV2144 | LV2065 | LV2031 | LV2167 | LV2128 | LV2194 | LV2140 | LV2075 | LV2131 | LV2192 | LV2165 | LV2087 | LV2064 | LV2117 | LV2098 | LV2100 | LV1785 | LV2091 | LV2095 | LV2187 | LV2190 | LV2186 | LV2077 | LV1760 | LV2123 | LV1780 | LV1789 | LV2179 | LV2173 | LV2166 | LV2157 | LV2082 | LV2072 | LV2058 | LV2161 | LV2093 | LV2052 | LV2078 | LV2174 | LV2164 | LV2030 | LV2163 | LV2066 | LV2067 | LV2094 | LV2106 | LV2145 | LV2069 | LV2171 | LV2147 | LV2159 | LV2162 | LV2054 | LV2047 | LV2139 | LV2120 | LV2071 | LV2175 | LV2034 | LV2081 | LV2152 | LV2160 | LV2089 | LV2178 | LV2096 | LV2130 | LV2076 | LV2172 | LV2168 | LV2185 | LV2158 | LV2088 | LV2181 | LV2118 | LV2195 | LV2124 | LV2074 | LV2092 | LV2073 | LV2191 | LV2188 | LV2184 | LV2137 | LV2057 | LV2127
700 | 1327 | 1929 | 1299 | 1058 | 1553 | 1300 | 993 | 1148 | 1844 | 1231 | 2234 | 1029 | out of range | 1179 | 1396 | 1082 | 1238 | Too low | 1031 | 1388 | 1036 | 971 | 1304 | 1662 | 1204 | 1051 | 1092 | out of range | 1336 | 1848 | 1401 | 1172 | 966 | 988 | 975 | 1036 | 1311 | 1264 | 1659 | 1148 | 1454 | 1848 | 1714 | 1881 | 1431 | 1607 | 2215 | 1306 | 1998 | 1737 | 1666 | 981 | 1208 | 1282 | 1202 | 1321 | 0 | 1293 | 1771 | 1709 | 1518 | 1441 | 1248 | 1163 | 1238 | 1495 | 1390 | 951 | 1307 | 1085 | 1258 | 1310 | 1118 | 1390 | 1322 | 1143 | 1739 | 1285 | 1292 | 1496 | 1613 | 1455 | 1365 | 1516 | 1296 | 1057 | 1631 | 1231 | 1327 | 1805 | 1138 | 896 | 1055 | 1031 | 1029 | 1086 | 1357 | 1113 | 1191 | 1802 | 1662 | 1635 | 1440 | 1294 | 1239 | 1688 | 961 | 1550 | 1667 | 1497 | 1719 | 1058 | 1207 | 1488 | 1452 | out of range | 1229 | 1172 | 1735 | 1662
750 | 1987 | 2784 | 1944 | 1666 | 2298 | 1963 | 1422 | 1763 | 2721 | 1858 | 3179 | 1598 | 1112 | 1785 | 1894 | 1711 | 1819 | 1335 | 1546 | 1930 | 1586 | 1535 | 1989 | 2435 | 1857 | 1583 | 1673 | out of range | 2044 | 2748 | 2078 | 1727 | 1445 | 1518 | 1455 | 1632 | 1983 | 1926 | 2341 | 1782 | 2176 | 2574 | 2446 | 2752 | 2066 | 2380 | 2706 | 1975 | 2843 | 2476 | 2471 | 1466 | 1857 | 1966 | 1813 | 1939 | 1457 | 1941 | 2569 | 2485 | 2226 | 2191 | 1887 | 1746 | 1857 | 2153 | 2028 | 1502 | 2002 | 1681 | 1953 | 1988 | 1749 | 1927 | 1984 | 1727 | 2515 | 1970 | 1989 | 2140 | 2436 | 2156 | 2070 | 2196 | 1922 | 1556 | 2378 | 1895 | 2006 | 2676 | 1753 | 1268 | 1634 | 1532 | 1581 | 1595 | 1888 | 1723 | 1807 | 2683 | 2324 | 2337 | 2100 | 1851 | 1749 | 2276 | 1578 | 2144 | 2275 | 2203 | 2483 | 1601 | 1835 | 2147 | 2008 | 1496 | 1858 | 1718 | 2572 | 2436
775 | 2365 | 3300 | 2313 | 2007 | 2745 | 2375 | 1762 | 2102 | 3233 | 2213 | 3754 | 1923 | 1358 | 2126 | 2172 | 2065 | 2123 | 1615 | 1846 | 2240 | 1891 | 1856 | 2391 | 2903 | 2223 | 1866 | 1995 | 1052 | 2451 | 3302 | 2471 | 2011 | 1762 | 1828 | 1760 | 1959 | 2374 | 2313 | 2753 | 2141 | 2613 | 3002 | 2888 | 3280 | 2432 | 2843 | 3271 | 2363 | 3354 | 2927 | 2962 | 1788 | 2219 | 2357 | 2152 | 2291 | 1793 | 2310 | 3060 | 2969 | 2655 | 2636 | 2252 | 2070 | 2210 | 2552 | 2402 | 1814 | 2410 | 2054 | 2364 | 2379 | 2092 | 2234 | 2365 | 2065 | 2982 | 2365 | 2392 | 2525 | 2941 | 2579 | 2477 | 2599 | 2275 | 1849 | 2875 | 2273 | 2404 | 3221 | 2103 | 1572 | 1963 | 1843 | 1880 | 1876 | 2185 | 2062 | 2151 | 3222 | 2716 | 2818 | 2471 | 2212 | 1985 | 2679 | 1850 | 2513 | 2684 | 2699 | 2951 | 1915 | 2187 | 2525 | 2337 | 1761 | 2208 | 2017 | 3085 | 2910
800 | 2791 | 3885 | 2737 | 2391 | 3269 | 2834 | 2120 | 2486 | 3850 | 2614 | 4391 | 2283 | 1614 | 2517 | 2474 | 2476 | 2465 | 1884 | 2177 | 2576 | 2224 | 2219 | 2840 | 3434 | 2639 | 2185 | 2359 | 1298 | 2922 | 3925 | 2913 | 2324 | 2100 | 2172 | 2086 | 2338 | 2810 | 2754 | 3211 | 2551 | 3103 | 3474 | 3376 | 3875 | 2851 | 3365 | 3916 | 2805 | 3921 | 3433 | 3506 | 2140 | 2636 | 2801 | 2539 | 2685 | 2172 | 2729 | 3619 | 3507 | 3148 | 3152 | 2660 | 2432 | 2610 | 2995 | 2828 | 2158 | 2864 | 2390 | 2820 | 2874 | 2484 | 2578 | 2797 | 2448 | 3510 | 2814 | 2845 | 2948 | 3520 | 3058 | 2953 | 3052 | 2681 | 2178 | 3381 | 2707 | 2853 | 3816 | 2498 | 1882 | 2337 | 2188 | 2214 | 2193 | 2522 | 2440 | 2544 | 3831 | 3160 | 3326 | 2844 | 2462 | 2281 | 3064 | 2150 | 2898 | 3115 | 3213 | 3480 | 2271 | 2593 | 2947 | 2700 | 2044 | 2598 | 2357 | 3675 | 3447
825 | 3261 | 4533 | 3213 | 2818 | 3841 | 3348 | 2517 | 2914 | 4528 | 3155 | 5098 | 2693 | 1758 | 2946 | 2801 | 2941 | 2831 | 2181 | 2555 | 2962 | 2599 | 2630 | 3338 | 4048 | 3106 | 2545 | 2774 | 1573 | 3451 | 4634 | 3404 | 2673 | 2480 | 2551 | 2457 | 2762 | 3306 | 3245 | 3726 | 3030 | 3647 | 3985 | 3924 | 4547 | 3321 | 3949 | 4644 | 3288 | 4551 | 3991 | 4112 | 2547 | 3104 | 3303 | 2965 | 3121 | 2615 | 3197 | 4242 | 4114 | 3693 | 3670 | 3126 | 2835 | 3056 | 3488 | 3264 | 2382 | 3370 | 2843 | 3334 | 3333 | 2926 | 2951 | 3280 | 2881 | 4095 | 3316 | 3356 | 3436 | 4173 | 3589 | 3471 | 3558 | 3124 | 2545 | 3932 | 3198 | 3357 | 4492 | 2943 | 2220 | 2756 | 2570 | 2573 | 2544 | 2882 | 2869 | 2989 | 4508 | 3651 | 3938 | 3356 | 2854 | 2640 | 3595 | 2570 | 3291 | 3588 | 3765 | 4084 | 2670 | 3072 | 3426 | 3065 | 2363 | 3042 | 2737 | 4328 | 4052
850 | 3800 | 5241 | 3749 | 3289 | 4491 | 3903 | 2962 | 3408 | 5240 | 3662 | 5868 | 3146 | 2132 | 3429 | 3160 | 3455 | 3240 | 2515 | 2970 | 3375 | 3018 | 3091 | 3902 | 4723 | 3632 | 2941 | 3233 | 1826 | 4043 | 5412 | 3948 | 3050 | 2913 | 2963 | 2867 | 3237 | 3861 | 3799 | 4283 | 3552 | 4249 | 4574 | 4526 | 5292 | 3837 | 4586 | 2266 | 3835 | 5257 | 4623 | 4789 | 3002 | 3625 | 3865 | 3441 | 3613 | 3116 | 3703 | 4944 | 4768 | 4300 | 4412 | 3644 | 3274 | 3545 | 4036 | 3763 | 2990 | 3975 | 3337 | 3903 | 3895 | 3436 | 3361 | 3820 | 3367 | 4764 | 3876 | 3924 | 3967 | 4883 | 4187 | 4066 | 4125 | 3627 | 2951 | 4557 | 3757 | 3918 | 5229 | 3427 | 2602 | 3226 | 2998 | 2988 | 2945 | 3284 | 3341 | 3483 | 5262 | 4185 | 4576 | 3951 | 3291 | 2922 | 4082 | 2938 | 3840 | 4111 | 4325 | 4750 | 3098 | 3562 | 3939 | 3513 | 2721 | 3526 | 3163 | 5063 | 4723
t1 | 231 | 219 | 217 | 251 | 201 | 215 | 167.8 | 238 | 215.2 | 192.3 | 228 | 237 | 265 | 238 | 310 | 220 | 329 | 369 | 231 | 258 | 267 | 225 | 227 | 196.5 | 232 | 272 | 237 | 224 | 213 | 202 | 234 | 353 | 198 | 251 | 214 | 233 | 222 | 217 | 224 | 214 | 219 | 241 | 224 | 207 | 224 | 220 | 191.5 | 229 | 220 | 209 | 218 | 181.6 | 230 | 222 | 249 | 243 | 183 | 243 | 197 | 210 | 207 | 180 | 232 | 270 | 240 | 220 | 252 | 193 | 211 | 248 | 226 | 218 | 237 | 261.5 | 227.5 | 219 | 207 | 225.8 | 228 | 222 | 199 | 209 | 212 | 218 | 235 | 228.3 | 231 | 214 | 221 | 212 | 241 | 175 | 228 | 213 | 285 | 251 | 268 | 256 | 234 | 207 | 228 | 186 | 186 | 294 | 384 | 201 | 368 | 219 | 218 | 247 | 199 | 237 | 232 | 246 | 254 | 249 | 257 | 251 | 196 | 201
A1 | 2693 | 3417 | 2495 | 2772 | 2663 | 2614 | 1287 | 2563.6 | 3563 | 1812 | 3944 | 2431 | 581 | 2588 | 2818 | 2493 | 3460 | 2958 | 2152 | 2576 | 2627 | 2261 | 2764 | 2725 | 2700 | 2630 | 2489 | 1949 | 2700 | 3284 | 2862 | 3625 | 1676 | 2442 | 1840 | 2497 | 2656 | 2555 | 2806 | 2454 | 2854 | 3082 | 2976 | 3239 | 2585 | 3105 | 3531 | 2732 | 3336 | 2765 | 3141 | 1559 | 2687 | 2712 | 2742 | 2676 | 2322 | 2899 | 2807 | 2995 | 2630 | 2125 | 2638 | 2872 | 2703 | 2602 | 2911 | 1367 | 2501 | 2519 | 2834 | 2648 | 2599 | 2572 | 2694 | 2274 | 2840 | 2771 | 2848 | 2586 | 2990 | 2619 | 2653 | 2659 | 2627 | 2056 | 3095 | 2506 | 2696 | 3356 | 2705 | 1166 | 2363 | 1907 | 2763 | 2247 | 2591 | 2804 | 2579 | 3278 | 2772 | 2402 | 2053 | 1493 | 1921 | 1755 | 2110 | 2051 | 2391 | 3385 | 2728 | 2379 | 2751 | 2982 | 2431 | 2482 | 2926 | 2447 | 2945 | 2797
x0 | 699 | 702.3 | 703 | 704 | 700 | 701.7 | 688.2 | 700 | 710.8 | 675 | 701 | 703 | 405 | 702 | 698.5 | 706 | 700 | 660.5 | 703 | 706 | 700.4 | 703 | 699 | 703.6 | 702.8 | 707 | 705 | 775 | 703 | 702 | 702 | 709 | 693 | 702.4 | 697 | 705 | 701.6 | 701 | 703 | 708 | 700 | 694 | 700 | 701 | 704 | 702 | 699 | 699 | 698 | 699 | 698 | 697.5 | 705 | 703 | 703 | 699 | 602.9 | 706 | 701 | 703 | 701 | 690 | 701 | 703 | 704 | 700 | 700 | 663 | 695 | 683 | 700 | 703 | 698.6 | 702.7 | 701.7 | 701.2 | 699.5 | 702.1 | 701.5 | 701.6 | 704 | 700 | 701 | 700 | 701 | 701.8 | 690 | 701 | 701 | 700 | 704.4 | 684 | 702.7 | 698 | 694.7 | 698 | 701 | 701.5 | 703 | 701.6 | 704.6 | 701 | 704.5 | 498 | 463.4 | 656.8 | 473 | 674.5 | 692 | 695.7 | 702 | 706 | 713 | 705 | 688 | 750 | 702 | 701 | 701 | 702
y0 | -1379 | -1457 | -1161 | -1670 | -1108 | -1299 | -402.4 | -1410 | -1550 | -831 | -1700 | -1368 | -1004 | -1386 | -1435 | -1338 | -2215 | -2427 | -1092 | -1128 | -1583 | -1259 | -1473 | -1015 | -1460 | -1505 | -1334 | -897 | -1315 | -1408 | -1437 | -2353 | -772 | -1433 | -893 | -1401 | -1323 | -1278 | -1115 | -1213 | -1397 | -1309 | -1255 | -1338 | -1098 | -1460 | -1331 | -1428 | -1361 | -1034 | -1496 | -604 | -1417 | -1390 | -1504 | -1362 | -828 | -1537 | -1022 | -1245 | -1101 | -799 | -1378 | -1674 | -1421 | -1103 | -1524 | -684 | -1254 | -1613 | -1567 | -1299 | -1492 | -1155 | -1351 | -1117 | -1108 | -1460 | -1537 | -1072 | -1315 | -1154 | -1264 | -1134 | -1315 | -982 | -1601 | -1260 | -1349 | -1544 | -1519 | -394 | -1278 | -900 | -1782 | -1176 | -1215 | -1672 | -1353 | -1453 | -1057 | -765 | -548 | -1661 | -2314 | -496 | -2930 | -750 | -823 | -1965 | -979 | -1257 | -1389 | -1431 | -1094 | -987 | -1666 | -1261 | -1184 | -1109
Fit | 758.1566 | 711.0938 | 761.1007 | 781.184 | 737.4251 | 758.1589 | 799.7763 | 774.7804 | 715.9915 | 767.4795 | 692.5197 | 787.2003 | 849.0595 | 772.893 | 768.7726 | 776.7087 | 772.6527 | 817.5265 | 794.0718 | 764.2112 | 790.6153 | 792.0631 | 757.277 | 729.884 | 766.95 | 792.7719 | 781.2772 | 871.3862 | 753.0392 | 715.3287 | 751.5514 | 781.6205 | 799.6425 | 795.1018 | 801.1133 | 783.7444 | 757.9205 | 761.5933 | 733.4791 | 772.2288 | 744.4973 | 718.3025 | 726.8513 | 713.3424 | 751.6672 | 732.0843 | 693.4983 | 757.5547 | 706.0927 | 725.1813 | 727.4912 | 797.5049 | 766.9121 | 758.9918 | 771.0658 | 761.5782 | 645.336 | 761.1108 | 721.9524 | 726.2097 | 741.6712 | 745.9069 | 765.1332 | 776.7442 | 767.4519 | 745.7176 | 755.2086 | 800.2751 | 756.9207 | 779.2186 | 758.2368 | 757.4295 | 775.289 | 764.2855 | 758.0374 | 777.173 | 724.7216 | 758.6747 | 757.2556 | 746.9435 | 730.4479 | 745.3723 | 751.3372 | 742.6778 | 762.6487 | 794.2196 | 731.3063 | 763.7562 | 755.4356 | 717.4544 | 774.5541 | 817.054 | 784.1762 | 794.5062 | 791.6095 | 792.6338 | 767.0389 | 777.4189 | 771.291 | 718.2757 | 734.2521 | 733.7856 | 751.8375 | 769.6258 | 782.8632 | 735.4921 | 792.6934 | 746.5491 | 735.7959 | 740.9158 | 726.0862 | 787.6151 | 768.135 | 746.5708 | 757.3347 | 804.3156 | 766.8609 | 780.6613 | 722.3549 | 729.6198
| S1T |  |  |  |  | S1B |  |  |  |  | S2T |  |  |  |  | S2B |  |  |  |  | S3T |  |  |  |  | S3B |  |  |  |  | S4T |  |  |  |  | S4B |  |  |  |  | S1L | [9 empty] | S1R | [9 empty] | S2L | [9 empty] | S2R | [9 empty] | S3L | [9 empty] | S3R | [9 empty] | S4L | [9 empty] | S4R
[67 empty] | -1103.46529 ± 117.68196 | [46 empty] | -1094.26309 ± 355.23043
[67 empty] | 700.4559 ± -- | [46 empty] | 688.21611 ± 3.48837E7
[67 empty] | 2602.22632 ± -- | [46 empty] | 2432.14869 ± 3.34402E8
[67 empty] | 219.60391 ± 6.89817 | [46 empty] | 253.71341 ± 25.81445

## Logs
Gas | Filled | Emptied | Check Up | Pressure (Torr) | Time | Date | PMT Warm Ups | Board | Start Time | End Time | Voltage | Date
|  |  | X | 727.48 | 12:34PM | 2024-03-08 |  | S3B | 12:30PM | 5PM | 800 | 2024-03-09
|  |  | X | 727.53 | 9:45AM | 2024-03-11 |  | S3B | 9AM | 4PM | 800 | 2024-03-11
| X |  |  | 727.14 | 4:30PM | 2024-03-11 |  | S1T | 8:15AM | 11:30AM | 800 | 2024-03-12
|  |  | X | 726.14 | 10:12AM | 2024-03-12 |  | S1T | 11:40AM | 3:45PM | 800 | 2024-03-12
|  | X | [5 empty] | S2T | 8AM | 10:45AM | 800 | 2024-03-13
| X |  |  | 726 | 5:10PM | 2024-03-12 |  | S2T | 11:40AM | 4PM | 800 | 2024-03-13
|  |  | X | 725.48 | 9:35AM | 2024-03-13 |  | S3T | 8:00AM | 12:30PM | 800 | 2024-03-14
| X |  |  | 726.8 | 11:30AM | 2024-03-13 |  | S3T | 12:40PM |  | 800 | 2024-03-14
|  |  | X | 725.44 | 9:30AM | 2024-03-14 |  | S4T | 8:10 |  | 800 | 2024-03-15
|  |  |  | 721.36 | 8:20 | 2024-03-15 |  | S4B | 7:45 |  | 800 | 2024-03-18
|  |  |  | 722.02 | 7:40 | 2024-03-18 |  | S1B | 8:05 |  | 800 | 2024-03-19
|  | x |  |  |  | 2024-03-18 |  | S2B | 8:00 |  | 800 | 2024-03-20
| x |  |  |  | 4:30pm | 2024-03-18 |  | S2B | 11:42AM | 4:20PM | 800 | 2024-03-25
|  |  |  | 719.09 | 8:05 | 2024-03-19 |  | S1T | 11:30 | 4:05PM | 800 | 2024-03-26
|  |  |  | 721 | 8:00 | 2024-03-20 |  | S1T | 11:39AM |  | 800 | 2024-03-27
| X |  |  | 724.53 | 11:45 | 2024-03-25
|  |  | X | 723.86 | 4:20PM | 2024-03-25
| X |  |  | 724.51 | 11:19AM | 2024-03-26
| X |  |  | 725.75 | 11:31AM | 2024-03-27
| x |  |  | 724.07 | 1:58PM | 2024-04-02

## Settings
LED Setting | Amplitude (V) | Offset (V) | Frequency (kHz) | Pulse Width (ns) | Rise Time (ns) | Fall Time (ns) | Delay (ns) | Load | Photoelectrons Equivalent
1 | 6.18 | 3.09 | 1 | 200 | 8.4 | 50 | 10 | HiZ | 130
2 | 6.3 | 3.15 | 1 | 200 | 8.4 | 50 | 10 | HiZ | 733
3 | 6.42 | 3.21 | 1 | 200 | 8.4 | 50 | 10 | HiZ | 1889
DAQ Config. | Signal Width | Input Amp. | Jumper Value | Integration Long | Threshold | Window Start | Window Width | Center (ch) | V (mV) | Charge per Channel |  | 50 pe is | 1154.0808
| 50 | 300 | 300 | 36 | 0x601 | 16784 | 1840 | 7204 | 500 | 6.9406e-15 | C/ch | 1500 | 34622.424
[12 empty] | 2500 | 57704.04
[12 empty] | 1949.6035 | 45000

## LED Position Correction
Board | PMT | -6 | -4 | -2 | 0 | 2 | 4 | 6 | mm | Center |  | PMT Warm up with light before scanning | Voltage | LED Setting | Problems? |  |  | Position | S1T5 | S1T5repeat | S1T4 | S1T3 | S1T2 | S1T1 | Home | -20.6
S2B5 | LV2201 | 43399 | 45150 | 46106 | 44869 | 40782 | 38551 | 33397 | ch | -2.5886 | mm | 1:26pm->1:34pm | 800 | 3 | *Peak still moved after turning PMT off, moving LED, PMT back on. *At 800 V starting at 10:26am. I did not restart the arduino, so I can only assume I am at -4 mm now after moving backward from -6 mm. If this measurement does not eclipse the -6 mm after 15 minutes of warming up with a light, then I will reposition the LED after restarting the arduino code. *peak was at 44854 which was smaller than the intensity further away, so I will redo measurement after reloading arduino. *RESTARTING* Loose connections are hopefully fixed for the motor and braids are added for better grounding between stand and DAQ. Started measurement with warm up at 800 V 1:26pm. Peak is more stable so far, but increasing. *Motor never moved because one phase was disconnected so redoing -4 mm measurement (previous was 42807 ch). * There is some confusion with the measurements. Removing 43114 at 2 because 40782 was for sure at 2 and 38551 at 4. |  |  | -6 | 47676 | 49355 | 17223 | 45364 | 41266 | 49158 | Run | 1 | -1 | 20.6
S2B4 | LV2032 | 39145 | 42489 | 44570 | 43547 | 43221 | 39242 | 34083 | ch | -0.96697 | mm |  |  |  | *when measuring 6mm the first time I think I might have gone too far so 8mm is probably 27063. |  |  | -4 | 53581 | 53823 | 18534 | 50492 | 48848 | 53348 | PMT5 | 0
S2B3 | LV2070 | 35401 | 37802 | 38737 | 39088 | 38707 | 36525 | 32499 | [6 empty] | *QDC2.011 does not work even after changing jumper chip. Doing measurement on D10 instead of C10. |  |  | -2 | 55739 | 56111 | 19580 | 52936 | 54035 | 54229 | PMT4 | 60
S2B2 | LV1788 | 37226 | 38494 | 38431 | 37955 | 36721 | 34056 | 30032 | [6 empty] | *0,2,4 and 6 have all had double peaks on spectcl but very close so just going with it |  |  | 0 | 52107 | 54972 | 19254 | 54140 | 52849 | 56464 | PMT3 | 120
S2B1 | LV2121 | 50269 | 51565 | 51245 | 50507 | 49057 | 45297 | 39084 | [9 empty] | 2 | 52024 | 52049 | 19568 | 52595 | 51330 | 55951
[10 empty] | -1.84 | [6 empty] | 1.5742 | 4 | 47804 | 48112 | 18280 | 51336 | 49605 | 50312
[17 empty] | 0.78712 | 6 | 39857 | 39897 | 16101 | 45339 | 44272 | 43365
Equation | y = Intercept + B1*x^1 + B2*x^2 | [16 empty] | Center: |  | -1.34 | -0.44 | 0.04204 | 0.23423 | -0.69069 | mm
Plot | S2B5 | S2B4 | S2B3 | S2B2 | S2B1 | [12 empty] | Average: | -0.43888 | [5 empty] | mm
Weight | No Weighting | [16 empty] | Deviation: |  | 205.3199 | 0.25428 | 109.5788 | 153.3695 | 57.3742 | %
Intercept | 44424.8571 | 44300.5714 | 39307.6667 | 38082.4286 | 50821 | [12 empty] | Peak Int. |  | 55813.9744 | 19703.099 | 54107.3932 | 53655.299 | 56433.0403 | ch
B1 | -866.5714 | -411.2321 | -201.6071 | -574.4286 | -862.125 | [12 empty] | Int. Deviation |  | 1.5085 | 2.2793 | 0.060263 | 1.5027 | 0.054861 | %
B2 | -167.1429 | -212.5625 | -146.381 | -121.9821 | -167.1696
Residual Sum of Squares | 2272400 | 1021710 | 616403.381 | 296618.4286 | 1637300
R-Square(COD) | 0.98166 | 0.98734 | 0.98185 | 0.99482 | 0.98663 | [12 empty] | Reset for Each PMT |  |  |  |  | at calibrated voltage
Adj. R-Square | 0.97249 | 0.98101 | 0.97278 | 0.99223 | 0.97994 | [10 empty] | 55288 |  | Position | S1T5 | S1T4 | S1T3 | S1T2 | S1T1
[18 empty] | -6 |  |  |  | 37631 | 41209
| -2.5923 | -0.96732 | -0.68864 | -2.3546 | -2.5786 |  | -1.8363 |  |  |  | -1.902 | [6 empty] | -4 |  |  |  | 39980 | 40484
| 45548.067 | 44499.4681 | 39377.0839 | 38758.6919 | 51932.5348 | [5 empty] | -0.951 | [6 empty] | -2 |  |  |  | 39793 | 45167
[11 empty] | $ | [6 empty] | 0 |  |  |  | 41498 | 43844
[9 empty] | 0 | [8 empty] | 2 |  |  |  | 40616 | 43303
[18 empty] | 4 |  |  |  | 40020 | 40229
|  | b | -866.5714 | -411.2321 | -201.6071 | -574.4286 | -862.125 | [10 empty] | 6 |  |  |  | 37365 | 44162
|  | a | -167.1429 | -212.5625 | -146.381 | -121.9821 | -167.1696
|  |  | FALSE | -0.96732 | -0.68864 | -2.3546 | -2.5786

## Sheet1
PMT | LED1 | LED1 voltage | LED2 | LED2 voltage | LED3 | LED3 voltage | LED1 adjusted | Correction voltage | LED2 adjusted |  | LED3 adjusted |  |  |  | Fit Accuracy (%) | Gain Accuracy (%) | Drift (%) | PMT Resolution (%) | ABS Gain | Stats | Value | Units
LV2143 | 2057 | 758.1566 | 13883 | 759.626 | 44719 | 778.4382 | 2057 | 0 | 13982 | 760.8 | 45059 | 2739 | 17846 | 50816 | -2.09 | -2.09 | -1.9 | 13.758 | 2.09 | Average Gain Accuracy: | 0.558 | %
LV2177 | 2073 | 711.0938 | 14028 | 708.5569 | 44904 | 754.238 | 2073 | 712 | 14028 | 0 | 44904 | 0 | 0 | 0 | -0.84 | -0.13 | -0.89 | 2.298 | 0.13 | Gain Accuracy StDev: | 0.801 | %
LV1762 | 2097 | 761.1007 | 13923 | 753.361 | 44696 | 773.0777 | 2097 | 0 | 13923 | 0 | 44986 | 2734 | 18509 | 51660 | -0.63 | 0.13 | 0.15 | 0.886 | 0.13 | # of Gains Within 1%: | 297 | Gains
LV2085 | 2117 | 781.184 | 14108 | 776.5196 | 45348 | 779.7565 | 2117 | 0 | 14108 | 0 | 44881 | 2407 | 16408 | 51765 | #DIV/0! |  | #DIV/0! | #DIV/0! | na | % of Gains Within 1%: | 82.5 | %
LV2138 | 2094 | 737.4251 | 14083 | 735.1453 | 45101 | 749.3545 | 2094 | 0 | 14083 | 0 | 45101 | 3284 | 21439 | 61105 | #DIV/0! |  | #DIV/0! | #DIV/0! | na | # of LED1 Gains Within 1%:
LV2125 | 2088 | 758.1589 | 13962 | 758.3424 | 45013 | 788.0184 | 2088 | 0 | 13962 | 0 | 45013 | 2924 | 18409 | 48519 | #DIV/0! |  | #DIV/0! | #DIV/0! | na | % of LED1 Gains Within 1%:
LV2148 | 2090 | 799.7763 | 14022 | 803.5296 | 44862 | 834.1423 | 2090 | 0 | 14022 | 0 | 44862 | 2076 | 13722 | 37279 | -0.14 | -0.14 | -0.11 | 13.639 | 0.14
LV2060 | 2097 | 774.7804 | 13933 | 776.3276 | 44807 | 798.4455 | 2097 | 0 | 13933 | 0 | 44807 | 2465 | 16147 | 45420 | -0.55 | -0.55 | -0.15 | 2.356 | 0.55
LV2029 | 2084 | 715.9915 | 13888 | 712.3241 | 45439 | 751.5919 | 2084 | 0 | 13888 | 0 | 45439 | 3882 | 24004 | 57642 | -0.68 | -0.03 | 0 | 0.85 | 0.03
LV2050 | 2161 | 767.4795 | 13928 | 767.8394 | 44726 | 783.5601 | 2161 | 762.5 | 13928 | 0 | 44726 | 2611 | 17272 | 50127 | 0.8 | 0.8 | 0.66 | 10.439 | 0.8
LV2119 | 2060 | 692.5197 | 13747 | 688.2432 | 44395 | 714.9528 | 2060 | 0 | 13747 | 0 | 44395 | 0 | 0 | 0 | 0.77 | 0.77 | 0.8 | 1.992 | 0.77
LV2086 | 2094 | 787.2003 | 13715 | 780.9475 | 44384 | 785.3268 | 2094 | 0 | 13715 | 0 | 44384 | 0 | 0 | 0 | 0.77 | -0.27 | 0.5 | 0.867 | 0.27
LV2193 | 2137 | 849 | 13961 | 823 | 44880 | 844.1 | 2137 | 0 | 13961 | 0 | 44880 | 1623 | 12355 | 36593 | -0.29 | -0.29 | 0.46 | 11.032 | 0.29
LV2182 | 2034 | 772.893 | 13886 | 770.6308 | 44762 | 783.102 | 2034 | 0 | 13886 | 0 | 45150 | 2450 | 16675 | 49551 | 0.59 | 0.59 | 0.62 | 2.038 | 0.59
LV2134 | 2023 | 768.7726 | 13751 | 745.3385 | 44575 | 760.1468 | 2040 | 770.5 | 13850 | 746.5 | 44973 | 2388 | 18169 | 52583 | 0.22 | 0.22 | 0.42 | 0.816 | 0.22
LV2201 | 2114 | 776.7 | 14036 | 774 | 45094 | 789.9 | 2114 | 0 | 14036 | 0 | 45094 | 2499 | 16526 | 47911 | -0.57 | -0.57 | 3.08 | 14.943 | 0.57 | Average Drift Accuracy:
LV2032 | 2097 | 772.7 | 13995 | 766.9 | 45059 | 791.8 | 2097 | 0 | 13995 | 0 | 45059 | 2472 | 16671 | 46781 | -0.27 | -0.27 | 1.01 | 2.399 | 0.27
LV2070 | 2097 | 817.5265 | 13975 | 806.8264 | 44913 | 831.3762 | 2097 | 0 | 13975 | 0 | 44913 | 1889 | 13489 | 38899 | 0.03 | 0.03 | 0.39 | 0.922 | 0.03
LV1788 | 2051 | 794.0718 | 13863 | 795.7624 | 44801 | 834.9239 | 2079 | 796 | 14015 | 797.5 | 44989 | 2137 | 14186 | 38214 | -0.48 | -0.48 | -2.12 | 13.349 | 0.48
LV2121 | 2055 | 764.2112 | 13821 | 759.7544 | 44533 | 773.6797 | 2055 | 0 | 13821 | 0 | 44998 | 2530 | 17129 | 50985 | 0.16 | 0.16 | 0.07 | 2.254 | 0.16
LV2122 | 2109 | 790.6153 | 14093 | 789.8228 | 45308 | 817.1954 | 2109 | 0 | 14093 | 0 | 44807 | 2238 | 14921 | 41588 | -0.31 | -0.31 | 0.38 | 0.889 | 0.31
LV2090 | 2094 | 792.0631 | 14013 | 783.7688 | 44889 | 787.1426 | 2094 | 0 | 14013 | 0 | 44889 | 2221 | 15582 | 48860 | -0.14 | -0.14 | -0.85 | 12.828 | 0.14
LV2062 | 2104 | 757.277 | 14009 | 748.7443 | 45609 | 754.5437 | 2104 | 0 | 14009 | 0 | 44912 | 2837 | 19710 | 60000 | -0.48 | -0.48 | -0.77 | 2.196 | 0.48
LV2063 | 2083 | 729.884 | 13908 | 723.2843 | 46767 | 753.5108 | 0 | 0 | 13908 | 0 | 46767 | 0 | 0 | 0 | -0.43 | -0.43 | -0.01 | 0.877 | 0.43
LV2126 | 2078 | 766.95 | 13920 | 763.0189 | 44797 | 775.0403 | 2078 | 0 | 13920 | 0 | 45195 | 2616 | 17546 | 52190 | -0.77 | -0.77 | 0.82 | 14.587 | 0.77
LV2136 | 2102 | 792.7719 | 13882 | 783.4809 | 44620 | 807.1484 | 2102 | 0 | 13882 | 0 | 44620 | 0 | 0 | 0 | -0.81 | -0.81 | -0.31 | 2.398 | 0.81
LV2176 | 2093 | 781.2772 | 14336 | 775.7228 | 44976 | 802.6656 | 2087 | 773.7 | 14161 | 773.7 | 44976 | 2488 | 16443 | 44939 | 0.97 | 0.52 | 0.21 | 0.819 | 0.52
LV2079 | 0 | 871.3862 | 14168 | 837.7995 | 45355 | 835.3674 | 0 | 0 | 14168 | 0 | 45016 | 1300 | 11245 | 37171 | 2.82 | -0.29 | -0.11 | 11.522 | 0.29
LV2083 | 2092 | 753.0392 | 14073 | 745.1849 | 46091 | 748.0759 | 2092 | 0 | 14073 | 0 | 45102 | 2945 | 20432 | 62932 | -0.52 | -0.52 | -0.19 | 2.362 | 0.52
LV2200 | 2150 | 715.3287 | 14244 | 707.9941 | 45694 | 713.3643 | 2150 | 0 | 14151 | 707 | 45059 | 4004 | 27280 | saturates | -0.61 | -0.61 | -0.07 | 0.917 | 0.61
LV2146 | 2024 | 751.5514 | 13469 | 749.5902 | 44467 | 774.0244 | 2070 | 753 | 13806 | 752.5 | 44949 | 2818 | 18312 | 51016 | -1.94 | -1.94 | #DIV/0! | 14.369 | 1.94
LV2155 | 2046 | 781.6205 | 13746 | 778.1767 | 44517 | 802.1179 | 2068 | 784 | 13846 | 779.5 | 44917 | 2280 | 15387 | 44057 | -1.84 | -1.84 | #DIV/0! | 2.546 | 1.84
LV2169 | 2076 | 799.6425 | 13927 | 788.7905 | 44812 | 805.8348 | 2076 | 0 | 13927 | 0 | 44812 | 2139 | 15053 | 43514 | -1.36 | -1.36 | #VALUE! | 0.896 | 1.36
LV2084 | 2122 | 795.1018 | 14066 | 793.2778 | 44877 | 797.6741 | 2122 | 0 | 14066 | 0 | 44877 | 2190 | 14675 | 45528 | -0.29 | -0.29 | #DIV/0! | 14.804 | 0.29
LV2199 | 2160 | 801.1133 | 14260 | 796.4997 | 45728 | 813.8474 | 2160 | 0 | 14260 | 0 | 44785 | 2141 | 14560 | 42260 | -2.08 | -2.08 | #DIV/0! | 2.596 | 2.08
LV2144 | 2225 | 790.1 | 13980 | 777.6 | 42982 | 783.7 | 2180 | 787.5 | 13980 | 0 | 44917 | 2381 | 16107 | 47452 | -1.39 | -1.39 | #DIV/0! | 0.989 | 1.39
LV2065 | 2106 | 757.9205 | 13986 | 752.5198 | 44501 | 756.2886 | 2106 | 0 | 13986 | 0 | 44501 | 2821 | 19094 | 58244 | 1.73 | 1.73 | 0.55 | 13.804 | 1.73
LV2031 | 2105 | 761.5933 | 13955 | 748.6059 | 45605 | 755.0015 | 2105 | 0 | 13955 | 0 | 45605 | 2760 | 19728 | 59217 | -0.28 | -0.28 | 0.62 | 2.536 | 0.28
LV2167 | 2073 | 733.4791 | 13891 | 730.4136 | 44382 | 750.1731 | 2073 | 0 | 13891 | 0 | 44560 | 3190 | 20497 | 56849 | -0.27 | -0.27 | 0.18 | 0.898 | 0.27
LV2128 | 2080 | 772.2288 | 13945 | 765.4371 | 44586 | 775.431 | 2080 | 0 | 13945 | 0 | 44586 | 2536 | 17465 | 51967 | -3.24 | -3.24 | -2.73 | 15.388 | 3.24
0 | [6 empty] | 463.5128 |  | 114.5743 |  | 378.105 |  |  |  | -0.82 | -0.82 | -0.29 | 2.535 | 0.82
[7 empty] | 22.072 |  | 0.81839 |  | 0.84023 |  |  |  | -0.53 | 0.33 | 0.09 | 0.943 | 0.33
[15 empty] | -3.81 | -2.94 | -3.6 | 13.594 | 2.94
[15 empty] | -1.81 | -1.08 | -1.57 | 2.363 | 1.08
[15 empty] | -0.95 | -0.06 | -1.16 | 0.866 | 0.06
[15 empty] | 0.66 | 0.66 | 0.92 | 11.495 | 0.66
[15 empty] | 0.26 | 0.26 | 0.38 | 2.123 | 0.26
[15 empty] | 0.21 | 0.21 | 0.33 | 0.869 | 0.21
[15 empty] | -0.14 | -0.14 | 0.28 | 12.637 | 0.14
[15 empty] | -0.04 | -0.04 | 0.27 | 2.208 | 0.04
[15 empty] | 0.13 | 0.13 | 0.18 | 0.848 | 0.13
[15 empty] | -0.14 | -0.14 | 0.26 | 11.111 | 0.14
[15 empty] | -0.18 | -0.18 | -0.16 | 2.054 | 0.18
[15 empty] | -0.19 | -0.19 | 0.06 | 0.799 | 0.19
[15 empty] | -2.39 | -1.01 | -1.87 | 14.432 | 1.01
[15 empty] | -0.99 | 0.11 | -0.96 | 2.417 | 0.11
[15 empty] | -0.44 | -0.02 | -0.54 | 0.866 | 0.02
[15 empty] | -2.19 | -2.19 | -1.82 | 13.382 | 2.19
[15 empty] | -1.3 | -1.3 | -1.33 | 2.33 | 1.3
[15 empty] | -1.05 | -0.004 | -0.63 | 0.934 | 0.004
[15 empty] | 0.43 | 0.43 | 0.63 | 11.238 | 0.43
[15 empty] | 0.66 | 0.66 | 0.61 | 2.086 | 0.66
[15 empty] | 0.68 | -0.43 | 0.79 | 0.821 | 0.43
[15 empty] | -0.29 | -0.29 | 0.09 | 11.939 | 0.29
[15 empty] | 0.09 | 0.09 | 0.32 | 0 | 0.09
[15 empty] | -0.25 | -0.25 | 0.42 | 0.902 | 0.25
[15 empty] | 0.19 | 0.19 | -0.11 | 11.977 | 0.19
[15 empty] | 0.06 | 0.06 | 0.31 | 2.184 | 0.06
[15 empty] | 1.34 | -0.2 | 0.43 | 0.881 | 0.2
[15 empty] | #DIV/0! | 0.94 | -0.24 | #DIV/0! | 0.94
[15 empty] | -0.16 | -0.16 | -0.34 | 1.917 | 0.16
[15 empty] | #DIV/0! |  | #DIV/0! | #DIV/0! | na
[15 empty] | -1.06 | -1.06 | -0.88 | 11.261 | 1.06
[15 empty] | -0.57 | -0.57 | -0.55 | 2.112 | 0.57
[15 empty] | -0.45 | 0.43 | -0.13 | 0.848 | 0.43
[15 empty] | 0.1 | 0.1 | #DIV/0! | 15.557 | 0.1
[15 empty] | -0.85 | -0.85 | #DIV/0! | 2.658 | 0.85
[15 empty] | -0.85 | -0.85 | #DIV/0! | 0.968 | 0.85
[15 empty] | -0.33 | -0.62 | 5.18 | 11.706 | 0.62
[15 empty] | 2.34 | 1.14 | 2.64 | 2.037 | 1.14
[15 empty] | -0.05 | -0.05 | 1.24 | 0.803 | 0.05
[15 empty] | #DIV/0! |  | 0.15 | #DIV/0! | na
[15 empty] | 1.19 | 1.19 | 1.03 | 2.209 | 1.19
[15 empty] | 0.78 | 0.04 | 0.51 | 0.891 | 0.04
[15 empty] | -0.38 | -0.38 | 0.78 | 13.48 | 0.38
[15 empty] | 0.52 | 0.52 | 0.5 | 2.395 | 0.52
[15 empty] | 2.37 | 0.23 | 0.13 | 0.946 | 0.23
[15 empty] | 2.33 | 2.33 | 1.97 | 14.326 | 2.33
[15 empty] | 1.71 | 1.07 | -2.51 | 2.499 | 1.07
[15 empty] | 1.52 | 0.13 | #VALUE! | 0 | 0.13
[15 empty] | -3.75 | -1.45 | -3.37 | 14.13 | 1.45
[15 empty] | -3.94 | -1.41 | -7.16 | 2.457 | 1.41
[15 empty] | -1.2 | -0.11 | -0.65 | 0.884 | 0.11
[15 empty] | -2.64 | -1.55 | -1.93 | 14.467 | 1.55
[15 empty] | -1.85 | -1.11 | -1.35 | 2.503 | 1.11
[15 empty] | -1.08 | -0.18 | -0.95 | 0.943 | 0.18
[15 empty] | -1.16 | -1.16 | 1.82 | 13.776 | 1.16
[15 empty] | -0.52 | -0.52 | 0.65 | 2.456 | 0.52
[15 empty] | -0.42 | -0.42 | -0.09 | 1.015 | 0.42
[15 empty] | 1.04 | 1.04 | 0.82 | 14.326 | 1.04
[15 empty] | 0.47 | 0.47 | 0.76 | 2.559 | 0.47
[15 empty] | -0.27 | -0.27 | 0.27 | 0.989 | 0.27
[15 empty] | 2.78 | 2.78 | 2.57 | 0 | 2.78
[15 empty] | 1.82 | 1.82 | 1.65 | 2.419 | 1.82
[15 empty] | 1.59 | -0.48 | 0.94 | 0.923 | 0.48
[15 empty] | 5.62 | 3.67 | 1.81 | 11.416 | 3.67
[15 empty] | -0.14 | -0.14 | 0.4 | 2.232 | 0.14
[15 empty] | -4.69 | -0.18 | -0.34 | 0.912 | 0.18
[15 empty] | 0.28 | 0.28 | 0.39 | 12.346 | 0.28
[15 empty] | -0.1 | -0.1 | 0.07 | 2.238 | 0.1
[15 empty] | -1.12 | -1.12 | 0.01 | 0.888 | 1.12
[15 empty] | 0.24 | 0.24 | 0.22 | 11.971 | 0.24
[15 empty] | -0.32 | -0.32 | 0.06 | 2.236 | 0.32
[15 empty] | 1.33 | 1.33 | -0.01 | 0.89 | 1.33
[15 empty] | -1.3 | -1.3 | -0.66 | 15.581 | 1.3
[15 empty] | -0.78 | -0.78 | -0.48 | 2.606 | 0.78
[15 empty] | -1.39 | -0.99 | -0.25 | 0.969 | 0.99
[15 empty] | -0.96 | -0.96 | -0.59 | 15.433 | 0.96
[15 empty] | -0.39 | -0.39 | 0.01 | 2.668 | 0.39
[15 empty] | -0.93 | -0.93 | 0.12 | 1.003 | 0.93
[15 empty] | -2.84 | -0.43 | -2.34 | 14.496 | 0.43
[15 empty] | -2.62 | 0.33 | -1.89 | 2.558 | 0.33
[15 empty] | -2.26 | 0.14 | -1.91 | 0.993 | 0.14
[15 empty] | -3.24 | -0.91 | -2.54 | 14.159 | na
[15 empty] | -1.97 | -1.97 | -1.32 | 2.294 | na
[15 empty] | 0.45 | 0.45 | -0.56 | 0.923 | na
[15 empty] | -3.3 | -1.11 | -2.18 | 12.838 | 1.11
[15 empty] | -1.52 | -1.52 | -1.24 | 2.219 | 1.52
[15 empty] | -0.88 | -0.07 | -0.65 | 0.843 | 0.07
[15 empty] | -0.77 | -0.77 | -0.41 | 13.724 | 0.77
[15 empty] | 0.24 | 0.24 | 0.6 | 2.437 | 0.24
[15 empty] | -0.08 | -0.08 | #VALUE! | 0.994 | 0.08
[15 empty] | -3.14 | -0.43 | -2.26 | 12.574 | 0.43
[15 empty] | -1.9 | -1.9 | -1.47 | 2.169 | 1.9
[15 empty] | -1.13 | 0.63 | -0.94 | 0.827 | 0.63
[15 empty] | -2.19 | 1.04 | -0.87 | 14.793 | 1.04
[15 empty] | -1.03 | -1.03 | -0.85 | 2.504 | 1.03
[15 empty] | -0.64 | -0.64 | -0.28 | 1.02 | 0.64
[15 empty] | -1.69 | -1.69 | -0.81 | 12.785 | 1.69
[15 empty] | #DIV/0! |  | -0.4 | #DIV/0! | 0
[15 empty] | #DIV/0! |  | -0.17 | #DIV/0! | 0
[15 empty] | -2.59 | 0.14 | -2.26 | 12.799 | 0.14
[15 empty] | -1.51 | -1.51 | -0.97 | 2.269 | 1.51
[15 empty] | -1.92 | 0.35 | -0.33 | 0.969 | 0.35
[15 empty] | -1.6 | -0.1 | -0.8 | 14.272 | 0.1
[15 empty] | -0.55 | -0.55 | -0.11 | 2.772 | 0.55
[15 empty] | -0.5 | -0.5 | 0.23 | 0.909 | 0.5
[15 empty] | -2.49 | -0.1 | -1.42 | 13.08 | 0.1
[15 empty] | -1.64 | 0.08 | -0.85 | 2.258 | 0.08
[15 empty] | -0.94 | -0.04 | -0.24 | 0.832 | 0.04
[15 empty] | -1.55 | -0.82 | 0.57 | 14.942 | 0.82
[15 empty] | -0.95 | -0.95 | -0.19 | 2.437 | 0.95
[15 empty] | -1.2 | -0.01 | -0.13 | 0.94 | 0.01
[15 empty] | 0.38 | 0.38 | 0.33 | 12.808 | 0.38
[15 empty] | 0.03 | 0.03 | 0.81 | 2.285 | 0.03
[15 empty] | -0.17 | -0.17 | 0.29 | 0.926 | 0.17
[15 empty] | -0.19 | -0.19 | -2.21 | 12.691 | 0.19
[15 empty] | 0.02 | 0.02 | -0.78 | 2.207 | 0.02
[15 empty] | -0.38 | -0.38 | 0.12 | 0.917 | 0.38
[15 empty] | 0.14 | 0.14 | -2.9 | 13.79 | 0.14
[15 empty] | -0.24 | -0.24 | -1.47 | 2.456 | 0.24
[15 empty] | -0.68 | -0.02 | -0.18 | 0.984 | 0.02
[15 empty] | -0.38 | -0.38 | -30.14 | 12.667 | 0.38
[15 empty] | -0.55 | -0.55 | -27.67 | 2.227 | 0.55
[15 empty] | -0.2 | -0.2 | -24.45 | 0.866 | 0.2
[15 empty] | -0.53 | -0.53 | -2.25 | 12.59 | 0.53
[15 empty] | -0.45 | -0.45 | -0.73 | 2.167 | 0.45
[15 empty] | -0.2 | -0.2 | -0.22 | 0.833 | 0.2
[15 empty] | #DIV/0! |  | -1.21 | #DIV/0! | na
[15 empty] | #DIV/0! |  | -0.37 | #DIV/0! | na
[15 empty] | 0.46 | 0.46 | #DIV/0! | 0.878 | 0.46
[15 empty] | 0.33 | 0.33 | 0.84 | 12.387 | 0.33
[15 empty] | -0.54 | -0.54 | 0.79 | 2.183 | 0.54
[15 empty] | -0.41 | -0.41 | 0.54 | 0.852 | 0.41
[15 empty] | -0.57 | -0.57 | 0.66 | 12.979 | 0.57
[15 empty] | -0.01 | -0.01 | 0.72 | 2.257 | 0.01
[15 empty] | -0.9 | -0.27 | 0.35 | 0.877 | 0.27
[15 empty] | -0.05 | -0.05 | 3.47 | 14.721 | 0.05
[15 empty] | -0.17 | -0.17 | 1.07 | 2.411 | 0.17
[15 empty] | -0.48 | -0.478 | 0.21 | 0.918 | 0.478
[15 empty] | 1.32 | 1.32 | 1.82 | 16.212 | 1.32
[15 empty] | 4.64 | 0.48 | 2.28 | 2.398 | 0.48
[15 empty] | 0.39 | 0.39 | #VALUE! | 0.996 | 0.39
[15 empty] | 0.57 | 0.57 | 0.72 | 12.879 | 0.57
[15 empty] | -0.03 | -0.03 | 1.68 | 2.294 | 0.03
[15 empty] | -0.15 | -0.15 | -0.03 | 0.966 | 0.15
[15 empty] | 0.57 | 0.57 | 0.75 | 12.5 | 0.57
[15 empty] | -0.14 | -0.14 | 0.58 | 2.232 | 0.14
[15 empty] | -0.44 | 0.22 | -0.09 | 0.882 | 0.22
[15 empty] | -0.33 | -0.33 | 0.16 | 13.712 | 0.33
[15 empty] | 0.2 | 0.2 | 0.36 | 2.317 | 0.2
[15 empty] | -0.23 | -0.23 | 0.26 | 0.878 | 0.23
[15 empty] | -0.38 | -0.38 | -0.58 | 12.428 | 0.38
[15 empty] | -0.39 | -0.39 | -0.52 | 2.18 | 0.39
[15 empty] | -0.17 | -0.17 | -0.04 | 0.817 | 0.17
[15 empty] | -0.48 | -0.48 | 0.37 | 35.598 | 0.48
[15 empty] | 0.14 | 0.14 | 0.09 | 2.161 | 0.14
[15 empty] | 0.02 | 0.02 | 0.1 | 0.835 | 0.02
[15 empty] | 0.19 | 0.19 | -0.35 | 13.688 | 0.19
[15 empty] | 2.44 | 0.43 | 0.3 | 2.279 | 0.43
[15 empty] | -0.2 | -0.2 | 0.19 | 0.882 | 0.2
[15 empty] | 2.78 | 0.94 | 1.51 | 12.083 | 0.94
[15 empty] | 0.41 | 0.41 | 1.31 | 2.269 | 0.41
[15 empty] | -0.54 | -0.22 | 0.58 | 0.885 | 0.22
[15 empty] | 0.9 | 0.9 | 1.48 | 12.836 | 0.9
[15 empty] | -0.29 | -0.29 | 0.86 | 2.314 | 0.29
[15 empty] | 0 | -0.0044 | 0.42 | 0.962 | 0.0044
[15 empty] | 1.69 | 1.69 | -2.93 | 13.998 | 1.69
[15 empty] | 2.24 | 0.42 | 0.26 | 2.465 | 0.42
[15 empty] | 0.57 | -0.13 | 4.27 | 1.016 | 0.13
[15 empty] | -0.67 | -0.67 | 3.66 | 14.286 | 0.67
[15 empty] | -0.52 | -0.52 | 1.29 | 2.463 | 0.52
[15 empty] | -0.06 | -0.06 | 0.31 | 1.005 | 0.06
[15 empty] | 0 | 0 | -3.75 | 12.619 | 0
[15 empty] | -0.4 | -0.4 | -1.44 | 2.252 | 0.4
[15 empty] | 0.65 | 0.6469 | -0.22 | 0.936 | 0.6469
[15 empty] | -1.3 | -1.3 | -3.07 | 12.639 | 1.3
[15 empty] | 0.47 | 0.47 | -0.69 | 2.197 | 0.47
[15 empty] | -0.04 | -0.04 | -0.47 | 0.876 | 0.04
[15 empty] | 0 | 0 | 2.75 | 14.048 | 0
[15 empty] | -0.68 | -0.68 | 3.07 | 2.359 | 0.68
[15 empty] | -0.58 | -0.58 | 3.36 | 0.934 | 0.58
[15 empty] | -0.29 | -0.29 | -1.67 | 12.56 | 0.29
[15 empty] | -0.57 | -0.57 | -0.6 | 2.191 | 0.57
[15 empty] | -0.5 | -0.5 | -0.1 | 0.869 | 0.5
[15 empty] | -0.48 | -0.48 | -1.92 | 12.823 | 0.48
[15 empty] | -0.26 | -0.26 | -1.13 | 2.184 | 0.26
[15 empty] | -0.12 | -0.12 | -0.24 | 0.899 | 0.12
[15 empty] | -0.24 | -0.24 | -0.89 | 14.129 | 0.24
[15 empty] | -0.16 | -0.16 | -0.29 | 2.289 | 0.16
[15 empty] | -0.32 | -0.32 | -0.06 | 0.852 | 0.32
[15 empty] | -0.33 | -0.33 | -4.22 | 12.422 | 0.33
[15 empty] | -0.32 | -0.32 | -2.02 | 2.214 | 0.32
[15 empty] | -0.8 | -0.8 | -1.13 | 0.93 | 0.8
[15 empty] | 0 | 0 | 0.39 | 12.857 | 0
[15 empty] | 3.49 | 1.26 | 0.25 | 2.165 | 1.26
[15 empty] | 0.32 | 0.32 | 0.4 | 0.862 | 0.32
[15 empty] | -0.29 | -0.29 | 0.2 | 14.327 | 0.29
[15 empty] | -3.83 | -1.38 | -0.23 | 3.73 | 1.38
[15 empty] | 1.27 | 0.03 | 0.05 | 0.911 | 0.03
[15 empty] | -1.2 | -1.2 | 0.09 | 14.602 | 1.2
[15 empty] | -0.87 | -0.87 | -0.31 | 2.479 | 0.87
[15 empty] | -0.3 | -0.3 | -0.43 | 0.949 | 0.3
[15 empty] | -0.72 | -0.72 | 0.75 | 12.998 | 0.72
[15 empty] | -0.32 | -0.32 | 0.23 | 2.229 | 0.32
[15 empty] | -0.31 | -0.31 | 0.31 | 0.878 | 0.31
[15 empty] | -0.77 | -0.77 | -0.07 | 12.62 | 0.77
[15 empty] | -1.8 | -1.8 | 0.05 | 2.211 | 1.8
[15 empty] | -1.13 | 0.06 | 0.08 | 0.858 | 0.06
[15 empty] | -0.67 | -0.67 | 0.29 | 13.95 | 0.67
[15 empty] | -1.27 | -1.27 | 0.08 | 2.344 | 1.27
[15 empty] | -0.65 | -0.65 | 0.18 | 0.868 | 0.65
[15 empty] | -0.48 | -0.48 | 2.08 | 12.632 | 0.48
[15 empty] | -0.91 | -0.91 | 1.06 | 2.184 | 0.91
[15 empty] | -0.58 | -0.58 | 0.44 | 0.847 | 0.58
[15 empty] | -0.91 | -0.91 | -4.97 | 12.926 | 0.91
[15 empty] | -0.26 | -0.26 | -1.64 | 2.156 | 0.26
[15 empty] | #DIV/0! |  | -0.73 | #DIV/0! | na
[15 empty] | 2.87 | 2.87 | -4.09 | 13.599 | 2.87
[15 empty] | -0.5 | -0.5 | 0.04 | 2.297 | 0.5
[15 empty] | -0.66 | -0.66 | 0.15 | 0.839 | 0.66
[15 empty] | -0.19 | -0.19 | -0.82 | 12.452 | 0.19
[15 empty] | -0.24 | -0.24 | -0.36 | 2.227 | 0.24
[15 empty] | -0.24 | -0.24 | -0.32 | 0.875 | 0.24
[15 empty] | -0.53 | -0.53 | -3.44 | 13.021 | 0.53
[15 empty] | -0.06 | -0.06 | -1.41 | 2.273 | 0.06
[15 empty] | 1.34 | 0.02 | -0.24 | 0.934 | 0.02
[15 empty] | -1.2 | -1.2 | 6.97 | 14.217 | 1.2
[15 empty] | -0.86 | -0.86 | 3.14 | 2.723 | 0.86
[15 empty] | -0.48 | -0.48 | 0.59 | 0.936 | 0.48
[15 empty] | -0.24 | -0.24 | -2.8 | 15.465 | 0.24
[15 empty] | -0.25 | -0.25 | -1.03 | 2.664 | 0.25
[15 empty] | 0.5 | 0.5 | -0.01 | 1.048 | 0.5
[15 empty] | 0.47 | 0.47 | -1.42 | 12.37 | 0.47
[15 empty] | 0.53 | 0.53 | -0.87 | 2.245 | 0.53
[15 empty] | #DIV/0! | -4.78 | -0.3 | #DIV/0! | 4.78
[15 empty] | -0.19 | -0.19 | -2.01 | 12.261 | 0.19
[15 empty] | 0.27 | 0.27 | -0.13 | 2.201 | 0.27
[15 empty] | -0.45 | -0.45 | 0.01 | 0.886 | 0.45
[15 empty] | -0.14 | -0.14 | -3.11 | 13.495 | 0.14
[15 empty] | -0.21 | -0.21 | -1.12 | 2.312 | 0.21
[15 empty] | -0.23 | -0.23 | -0.01 | 0.931 | 0.23
[15 empty] | -0.48 | -0.48 | -2.12 | 12.344 | 0.48
[15 empty] | -0.8 | -0.8 | -1.11 | 2.167 | 0.8
[15 empty] | -0.1 | -0.1 | -0.15 | 0.914 | 0.1
[15 empty] | -0.72 | -0.72 | -0.83 | 13.189 | 0.72
[15 empty] | -1.59 | -1.59 | -1.21 | 2.351 | 1.59
[15 empty] | -0.48 | -0.48 | -0.56 | 0.882 | 0.48
[15 empty] | -0.86 | -0.86 | -3.74 | 14.409 | 0.86
[15 empty] | -0.92 | -0.92 | -1.63 | 2.494 | 0.92
[15 empty] | -0.08 | -0.08 | -0.18 | 0.941 | 0.08
[15 empty] | 0.14 | 0.14 | -1.79 | 13.029 | 0.14
[15 empty] | -0.56 | -0.56 | -0.48 | 2.313 | 0.56
[15 empty] | -0.81 | -0.1 | -0.07 | 0.95 | 0.1
[15 empty] | 0.05 | 0.05 | 1.47 | 13.517 | 0.05
[15 empty] | -0.06 | -0.06 | 0.48 | 2.387 | 0.06
[15 empty] | -0.05 | -0.05 | -0.1 | 0.9 | 0.05
[15 empty] | 0.24 | 0.24 | 2 | 15.012 | 0.24
[15 empty] | -0.24 | -0.24 | 0.64 | 2.563 | 0.24
[15 empty] | -0.15 | -0.15 | 0.21 | 0.946 | 0.15
[15 empty] | -0.33 | -0.33 | 1.06 | 13.712 | 0.33
[15 empty] | -0.5 | -0.5 | -0.9 | 2.441 | 0.5
[15 empty] | -0.18 | -0.18 | -1.29 | 0.964 | 0.18
[15 empty] | -1.65 | 0.76 | 1.04 | 0 | 0.76
[15 empty] | -0.29 | -0.29 | -0.89 | 2.221 | 0.29
[15 empty] | -0.16 | -0.16 | -0.56 | 0.824 | 0.16
[15 empty] | 0.9 | -0.57 | 2.6 | 10.618 | 0.57
[15 empty] | -0.25 | -0.25 | -0.29 | 2.148 | 0.25
[15 empty] | -0.6 | -0.6 | -0.48 | 0.908 | 0.6
[15 empty] | -0.43 | -0.43 | 0.81 | 15.686 | 0.43
[15 empty] | -0.78 | -0.78 | -1.12 | 2.368 | 0.78
[15 empty] | -0.27 | -0.27 | -0.52 | 0.958 | 0.27
[15 empty] | -0.77 | -0.77 | -12.59 | 11.996 | 0.77
[15 empty] | 0.2 | 0.2 | -13.42 | 2.153 | 0.2
[15 empty] | -0.01 | -0.01 | -12 | 0.962 | 0.01
[15 empty] | 0.05 | 0.05 | -0.62 | 14.231 | 0.05
[15 empty] | 0 | 0 | 0.09 | 2.15 | 0
[15 empty] | -0.24 | -0.24 | 0.06 | 0.853 | 0.24
[15 empty] | -4.32 | -4.32 | 2.27 | 15.077 | 4.32
[15 empty] | -0.6 | -0.6 | -0.14 | 2.436 | 0.6
[15 empty] | -0.27 | -0.27 | 0.09 | 0.9 | 0.27
[15 empty] | -1.6 | 0.57 | 1.33 | 13.256 | 0.57
[15 empty] | -0.24 | -0.24 | 0.36 | 2.305 | 0.24
[15 empty] | 0.03 | 0.03 | -0.09 | 0.82 | 0.03
[15 empty] | -0.82 | -0.82 | 0.92 | 14.642 | 0.82
[15 empty] | -0.04 | -0.04 | 0.15 | 2.051 | 0.04
[15 empty] | -0.36 | -0.36 | 0.16 | 0.879 | 0.36
[15 empty] | -1.35 | -1.35 | 0.8 | 15.106 | 1.35
[15 empty] | -0.44 | -0.44 | 0.39 | 2.447 | 0.44
[15 empty] | -1.16 | 0.33 | -0.12 | 0.892 | 0.33
[15 empty] | -0.77 | -0.77 | -0.55 | 14.443 | 0.77
[15 empty] | -0.51 | -0.51 | -0.81 | 2.462 | 0.51
[15 empty] | -0.2 | -0.2 | -0.18 | 0.937 | 0.2
[15 empty] | 0 | 0 | 0.87 | 12.571 | 0
[15 empty] | 0.36 | 0.36 | -0.48 | 2.228 | 0.36
[15 empty] | -0.36 | -0.15 | -0.26 | 0.901 | 0.15
[15 empty] | -0.14 | -0.14 | -1.77 | 12.494 | 0.14
[15 empty] | -0.29 | -0.29 | -0.45 | 2.221 | 0.29
[15 empty] | -0.1 | -0.1 | 0.05 | 0.905 | 0.1
[15 empty] | -0.19 | -0.19 | -0.44 | 13.693 | 0.19
[15 empty] | 0.16 | 0.16 | -0.18 | 2.268 | 0.16
[15 empty] | 0.06 | 0.06 | -0.02 | 0.864 | 0.06
[15 empty] | -0.62 | -0.62 | 0.37 | 12.314 | 0.62
[15 empty] | -0.54 | -0.54 | -0.18 | 2.14 | 0.54
[15 empty] | -0.04 | -0.04 | -0.06 | 0.878 | 0.04
[15 empty] | -0.33 | -0.33 | -1.89 | 12.183 | 0.33
[15 empty] | -0.26 | -0.26 | -1.04 | 2.127 | 0.26
[15 empty] | -0.72 | 0.16 | -0.25 | 0.844 | 0.16
[15 empty] | -0.86 | -0.86 | -1.52 | 13.689 | 0.86
[15 empty] | -0.94 | -0.94 | -0.65 | 2.3 | 0.94
[15 empty] | -1.03 | -0.2 | 0.14 | 0.851 | 0.2
[15 empty] | -0.62 | -0.62 | -1.07 | 12.506 | 0.62
[15 empty] | -0.56 | -0.56 | -0.07 | 2.169 | 0.56
[15 empty] | -0.44 | -0.44 | 0.34 | 0.819 | 0.44
[15 empty] | -1.06 | -1.06 | -2.05 | 13.041 | 1.06
[15 empty] | -0.31 | -0.31 | 0.1 | 2.228 | 0.31
[15 empty] | 3.2 | 0.4 | 0.11 | 0.852 | 0.4
[15 empty] | -0.91 | -0.91 | 1.35 | 14.368 | 0.91
[15 empty] | -0.89 | -0.89 | 0.31 | 2.429 | 0.89
[15 empty] | -0.27 | -0.27 | 0.33 | 0.925 | 0.27
Note on Q1: Dziubinski, Sean: % error = (experimental - actual) / actual *100